## Cal Certificate
Calibration Certificate
[10 empty] | Certificate No.: |  |  |  | 25-5465-09DBN Rev.00
Customer : | [6 empty] | Location of Calibration :
| Company 1 | [6 empty] | Company 1
| comp1 add | [6 empty] | comp1 add
Manufacturer: |  |  | AA |  |  |  | Description: |  |  |  |  | AA
Model No: |  |  | AA |  |  |  | Options Installed: |  |  |  |  | adads
Serial No: |  |  | 31214 |  |  |  | Customer Asset No: |  |  |  |  | adads
Calibration Date: |  |  | 2025-04-09 |  |  |  | Work Order No.: |  |  |  |  | 5465-09
Received Date: |  |  | 2025-04-09 |  |  |  | Customer P.O.: |  |  |  |  | adads
Temperature: |  |  | 5.5 °C |  |  |  | Humidity: |  |  |  |  | 7 %rh
Cal Method / Procedure :
As per SONJU Engineering Services Work Instruction 7W-02-54 Digital Multimeter Calibration Procedure. This is to certify that the above Measuring and Test Equipment has been calibrated by SONJU Engineering Services in accordance with relevant calibration procedure defined for this instrument.
This is to certify that the above Measuring and Test Equipment has been calibrated by SONJU Engineering Services in accordance with relevant calibration procedure defined for this instrument.
Uncertainty:
The expanded uncertainty of measurement for the results given in this certificate was estimated with a coverage factor of k = 2 at 95% level of confidence. The uncertainties of measurement were calculated using a budget worksheet based on JCGM 100:2008 Evaluation of measurement data - Guide to the Expression of Uncertainty in Measurement, 1st edition- 2008.
Traceability:
The instrument was calibrated using reference standard/s traceable to SI Units as maintained by NML-ITDI (Philippines) through SONJU Engineering Services (Philippines)
Reference Standard Used:
| Description |  |  |  | Serial No. |  |  | Cal Report Ref. | [5 empty] | Due Date
| adads |  |  |  | adads |  |  | adads | [5 empty] | Thursday, 4 September 2025
Remarks:
1. The results were the average of the three (3) readings and were obtained during the time of calibration and refer only to the particular instrument submitted. Should the instrument be modified or damaged in any way, or develop inconsistent readings, the results may not be valid and the unit will require recalibration.
2. The user should determine the suitability of this instrument for its intended use.
3. Preliminary Inspection: No response at high and low current.
4. A positve error is to subtracted to the UUT indication and negative error added to give true value.
5. As per customer request the next calibration due date is on: | [6 empty] | 2026-04-09
|  | DANICA BEA NERI | [8 empty] | Jennefer S. Ilas
|  | Calibration Officer | [8 empty] | Approved Signatory
Issue Date :
Measurement Results:
DC Voltage Test
|  | Range |  | Nominal Value |  | Measured Value |  |  | Error |  |  |  | Uncertainty
|  | 600 mV |  | 60 mV |  | 60.0 mV |  |  | 0.0 mV |  |  |  | 0.058 mV
|  |  |  | 540 mV |  | 540.0 mV |  |  | 0.0 mV |  |  |  | 0.058 mV
|  |  |  | -540 mV |  | -540.0 mV |  |  | 0.0 mV |  |  |  | 0.058 mV
|  | 6 V |  | 0.6 V |  | 0.600 V |  |  | 0.000 V |  |  |  | 0.000 58 V
|  |  |  | 5.4 V |  | 5.400 V |  |  | 0.000 V |  |  |  | 0.000 58 V
|  |  |  | -5.4 V |  | -5.400 V |  |  | 0.000 V |  |  |  | 0.000 58 V
|  | 60 V |  | 6 V |  | 6.00 V |  |  | 0.00 V |  |  |  | 0.005 8 V
|  |  |  | 30 V |  | 30.00 V |  |  | 0.00 V |  |  |  | 0.005 8 V
|  |  |  | 54 V |  | 54.00 V |  |  | 0.00 V |  |  |  | 0.005 8 V
|  |  |  | -6 V |  | -6.00 V |  |  | 0.00 V |  |  |  | 0.005 8 V
|  |  |  | -54 V |  | -54.00 V |  |  | 0.00 V |  |  |  | 0.005 8 V
|  | 600 V |  | 60 V |  | 60.0 V |  |  | 0.0 V |  |  |  | 0.058 V
|  |  |  | 540 V |  | 540.0 V |  |  | 0.0 V |  |  |  | 0.058 V
|  |  |  | -540 V |  | -540.0 V |  |  | 0.0 V |  |  |  | 0.058 V
|  | 1000 V |  | 100 V |  | 100 V |  |  | 0 V |  |  |  | 0.58 V
|  |  |  | 900 V |  | 900 V |  |  | 0 V |  |  |  | 0.58 V
|  |  |  | -900 V |  | -900 V |  |  | 0 V |  |  |  | 0.58 V
Resistance Test
|  | Range |  | Nominal Value |  | Measured Value |  |  | Error |  |  |  | Uncertainty
|  | 600 Ω |  | 0 Ω |  | 0.0 Ω |  |  | 0.0 Ω |  |  |  | 0.058 Ω
|  |  |  | 540 Ω |  | 540.0 Ω |  |  | 0.0 Ω |  |  |  | 0.058 Ω
|  | 6 kΩ |  | 5.4 kΩ |  | 5.400 kΩ |  |  | 0.000 kΩ |  |  |  | 0.000 58 kΩ
|  | 60 kΩ |  | 54 kΩ |  | 54.00 kΩ |  |  | 0.00 kΩ |  |  |  | 0.005 8 kΩ
|  | 600 kΩ |  | 540 kΩ |  | 540.0 kΩ |  |  | 0.0 kΩ |  |  |  | 0.058 kΩ
|  | 6 MΩ |  | 5.4 MΩ |  | 5.400 MΩ |  |  | 0.000 MΩ |  |  |  | 0.000 75 MΩ
|  | 50 MΩ |  | 45 MΩ |  | 45.00 MΩ |  |  | 0.00 MΩ |  |  |  | 0.015 MΩ
DANICA BEA NERI
Calibration Officer
Measurement Results:
AC Voltage Test
|  | Range |  | Nominal Value |  | Measured Value |  |  | Error |  |  |  | Uncertainty
|  | 600 mV |  | 60 mV / 50 Hz |  | 60.0 mV |  |  | 0.0 mV |  |  |  | 0.091 mV
|  |  |  | 60 mV / 1 kHz |  | 60.0 mV |  |  | 0.0 mV |  |  |  | 0.091 mV
|  |  |  | 540 mV / 50 Hz |  | 540.0 mV |  |  | 0.0 mV |  |  |  | 0.91 mV
|  |  |  | 540 mV / 1 kHz |  | 540.0 mV |  |  | 0.0 mV |  |  |  | 0.91 mV
|  | 6 V |  | 0.6 V / 50 Hz |  | 0.600 V |  |  | 0.000 V |  |  |  | 0.001 1 V
|  |  |  | 0.6 V / 1 kHz |  | 0.600 V |  |  | 0.000 V |  |  |  | 0.001 1 V
|  |  |  | 5.4 V / 50 Hz |  | 5.400 V |  |  | 0.000 V |  |  |  | 0.007 3 V
|  |  |  | 5.4 V / 1 kHz |  | 5.400 V |  |  | 0.000 V |  |  |  | 0.007 3 V
|  | 60 V |  | 6 V / 50 Hz |  | 6.00 V |  |  | 0.00 V |  |  |  | 0.009 3 V
|  |  |  | 6 V / 1 kHz |  | 6.00 V |  |  | 0.00 V |  |  |  | 0.009 3 V
|  |  |  | 30 V / 50 Hz |  | 30.00 V |  |  | 0.00 V |  |  |  | 0.009 3 V
|  |  |  | 30 V / 1 kHz |  | 30.00 V |  |  | 0.00 V |  |  |  | 0.009 3 V
|  |  |  | 54 V / 50 Hz |  | 54.00 V |  |  | 0.00 V |  |  |  | 0.071 V
|  |  |  | 54 V / 1 kHz |  | 54.00 V |  |  | 0.00 V |  |  |  | 0.071 V
|  | 600 V |  | 60 V / 50 Hz |  | 60.0 V |  |  | 0.0 V |  |  |  | 0.092 V
|  |  |  | 60 V / 1 kHz |  | 60.0 V |  |  | 0.0 V |  |  |  | 0.092 V
|  |  |  | 540 V / 50 Hz |  | 540.0 V |  |  | 0.0 V |  |  |  | 0.092 V
|  |  |  | 540 V / 1 kHz |  | 540.0 V |  |  | 0.0 V |  |  |  | 0.092 V
|  | 1000 V |  | 100 V / 50 Hz |  | 100 V |  |  | 0 V |  |  |  | 0.58 V
|  |  |  | 100 V / 1 kHz |  | 100 V |  |  | 0 V |  |  |  | 0.58 V
|  |  |  | 900 V / 50 Hz |  | 900 V |  |  | 0 V |  |  |  | 0.58 V
|  |  |  | 900 V / 1 kHz |  | 900 V |  |  | 0 V |  |  |  | 0.58 V
DC Current Test
|  | Range |  | Nominal Value |  | Measured Value |  |  | Error |  |  |  | Uncertainty
|  | 600 µA |  | 540 µA |  | 540.0 µA |  |  | 0.0 µA |  |  |  | 0.11 µA
|  | 6000 µA |  | 5 400 µA |  | 5 400 µA |  |  | 0 µA |  |  |  | 0.69 µA
|  | 60 mA |  | 54 mA |  | 54.00 mA |  |  | 0.00 mA |  |  |  | 0.005 9 mA
|  |  |  | -54 mA |  | -54.00 mA |  |  | 0.00 mA |  |  |  | 0.006 mA
|  | 400 mA |  | 360 mA |  | 360.0 mA |  |  | 0.0 mA |  |  |  | 0.082 mA
|  | 6 A |  | 3 A |  | 3.000 A |  |  | 0.000 A |  |  |  | 0.001 6 A
|  |  |  | 5.4 A |  | 5.400 A |  |  | 0.000 A |  |  |  | 0.001 6 A
|  | 10 A |  | 5 A |  | 5.00 A |  |  | 0.00 A |  |  |  | 0.005 9 A
|  |  |  | 9 A |  | 9.00 A |  |  | 0.00 A |  |  |  | 0.005 9 A
DANICA BEA NERI
Calibration Officer
Measurement Results:
AC Current Test
|  | Range |  | Nominal Value |  | Measured Value |  |  | Error |  |  |  | Uncertainty
|  | 600 µA |  | 540 µA / 50 Hz |  | 540.0 µA |  |  | 0.0 µA |  |  |  | 1.7 µA
|  |  |  | 540 µA / 1 kHz |  | 540.0 µA |  |  | 0.0 µA |  |  |  | 1.7 µA
|  | 6000 µA |  | 5 400 µA / 50 Hz |  | 5 400 µA |  |  | 0 µA |  |  |  | 13 µA
|  |  |  | 5 400 µA / 1 kHz |  | 5 400 µA |  |  | 0 µA |  |  |  | 13 µA
|  | 60 mA |  | 54 mA / 50 Hz |  | 54.00 mA |  |  | 0.00 mA |  |  |  | 0.22 mA
|  |  |  | 54 mA / 1 kHz |  | 54.00 mA |  |  | 0.00 mA |  |  |  | 0.22 mA
|  | 400 mA |  | 360 mA / 50 Hz |  | 360.0 mA |  |  | 0.0 mA |  |  |  | 0.72 mA
|  |  |  | 360 mA / 1 kHz |  | 360.0 mA |  |  | 0.0 mA |  |  |  | 0.72 mA
|  | 6 A |  | 5.4 A / 50 Hz |  | 5.400 A |  |  | 0.000 A |  |  |  | 0.010 A
|  |  |  | 5.4 A / 1 kHz |  | 5.400 A |  |  | 0.000 A |  |  |  | 0.011 A
|  | 10 A |  | 9 A / 50 Hz |  | 9.00 A |  |  | 0.00 A |  |  |  | 0.012 A
|  |  |  | 9 A / 1 kHz |  | 9.00 A |  |  | 0.00 A |  |  |  | 0.012 A
DANICA BEA NERI
Calibration Officer
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049) 513-5233/591-0807;
(02) 782-2153; 09236581107
Email: admin@sonjuengineering.com
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049) 513-5233/591-0807;
(02) 782-2153; 09236581107
Email: admin@sonjuengineering.com
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049) 513-5233/591-0807;
(02) 782-2153; 09236581107
Email: admin@sonjuengineering.com
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049) 513-5233/591-0807;
(02) 782-2153; 09236581107
Email: admin@sonjuengineering.com
Website: www.sonjuengineering.com

## DataSheet
CALIBRATION DATASHEET
|  | Work Order No.: | [5 empty] | 5465-09 | [13 empty] | Technical ID: |  |  |  | DBN
Unit Under Calibration
Description | [5 empty] | : | AA | [9 empty] | Received Date | [5 empty] | : | 2025-04-09
Manufacturer | [5 empty] | : | AA | [9 empty] | Calibration Date | [5 empty] | : | 2025-04-09
Model | [5 empty] | : | AA | [9 empty] | Options Installed | [5 empty] | : | adads
Serial Number | [5 empty] | : | 31214 | [9 empty] | Customer P.O. | [5 empty] | : | adads
Range | [5 empty] | : | See Specification Sheet | [9 empty] | Asset Number | [5 empty] | : | adads
Res/Readability | [5 empty] | : | See Specification Sheet | [9 empty] | Accuracy | [5 empty] | : | See Specification Sheet
Prev. SES Cal Cert: | [5 empty] | : | NA | [9 empty] | Previous Technician | [5 empty] | : | NA
Customer Details | [16 empty] | Calibration Location
Company |  |  | : | Company 1 | [15 empty] | In-House (SONJU Laboratory)
Address |  |  | : | comp1 add | [14 empty] | x | On-Site Calibration (Please Specify)
comp1 add
Reference Standard/s Used
Description | [10 empty] | Serial No. | [5 empty] | Cal Report Ref. | [9 empty] | Due Date
adads | [10 empty] | adads | [5 empty] | adads | [9 empty] | Thursday, 4 September 2025
adads | [10 empty] | adads | [5 empty] | adads | [9 empty] | Thursday, 4 September 2025
Accessory Used
Description | [10 empty] | Serial No. | [5 empty] | Brand | [9 empty] | Model
adads | [10 empty] | adads | [5 empty] | adads | [9 empty] | adads
adads | [10 empty] | adads | [5 empty] | adads | [9 empty] | adads
Environmental Condition | [16 empty] | Calibration Method
[5 empty] | Temperature | [5 empty] | Relative Humidity | [5 empty] | As per SONJU Engineering Services Work Instruction 7W-02-54 Digital Multimeter Calibration Procedure. This is to certify that the above Measuring and Test Equipment has been calibrated by SONJU Engineering Services in accordance with relevant calibration procedure defined for this instrument.
[6 empty] | Start: | 1 |  | oC |  |  | Start: | 2 |  | %rh
[6 empty] | End: | 10 |  | oC |  |  | End: | 12 |  | %rh
Observations
| - | [27 empty] | Adjustment
| - | [27 empty] | Needed?
| - | [28 empty] | YES
| - | [28 empty] | NO
Measurement Results
|  | DC Voltage Test | [28 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
|  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  |  | Measured Value | [8 empty] | Average |  |  | Error |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  | 600 mV | [5 empty] | 60 |  | mV | [5 empty] | 60 |  | 60 |  |  | 60 |  |  |  | 60 |  |  | 0 |  |  | 0.16 | 60.16 | 59.84 | PASS | CoC | 0.00077 | mV | 2 | 0.000385 | Annual Drift | 0 | mV | 1.7321 | 0 | Readability | 0.1 | mV | 3.4641 | 0.028868 | Repeatability | 0 | mV | 1.7321 | 0 | 0.02887 | 200.0711 | 2 | 0.05774 | mV | 0.0016 | mV | 0.05774 | unit
[8 empty] | 540 |  | mV |  | - |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 0.64 | 540.64 | 539.36 | PASS | CoC | 0.0021 | mV | 2 | 0.00105 | Annual Drift | 0 | mV | 1.7321 | 0 | Readability | 0.1 | mV | 3.4641 | 0.028868 | Repeatability | 0 | mV | 1.7321 | 0 | 0.028887 | 200.5292 | 2 | 0.057773 | mV | 0.0018 | mV | 0.057773 | mV
[8 empty] | -540 |  | mV |  | - |  |  |  | -540 |  | -540 |  |  | -540 |  |  |  | -540 |  |  | 0 |  |  | 0.64 | -539.36 | -540.64 | PASS | CoC | 0.0021 | mV | 2 | 0.00105 | Annual Drift | 0 | mV | 1.7321 | 0 | Readability | 0.1 | mV | 3.4641 | 0.028868 | Repeatability | 0 | mV | 1.7321 | 0 | 0.028887 | 200.5292 | 2 | 0.057773 | mV | 0.0018 | mV | 0.057773 | mV
|  | 6 V | [5 empty] | 0.6 |  | V |  | - |  |  |  | 0.6 |  | 0.6 |  |  | 0.6 |  |  |  | 0.6 |  |  | 0 |  |  | 0.0016 | 0.6016 | 0.5984 | PASS | CoC | 2.1e-06 | V | 2 | 1.05e-06 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.00028868 | 200.0053 | 2 | 0.00057735 | V | 1.8e-06 | V | 0.00057735 | V
[8 empty] | 5.4 |  | V |  | - |  |  |  | 5.4 |  | 5.4 |  |  | 5.4 |  |  |  | 5.4 |  |  | 0 |  |  | 0.0064 | 5.4064 | 5.3936 | PASS | CoC | 1.8e-05 | V | 2 | 9e-06 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | 1.0878e-15 | V | 1.7321 | 6.2804e-16 | 0.00028882 | 200.3888 | 2 | 0.00057763 | V | 1e-05 | V | 0.00057763 | V
[8 empty] | -5.4 |  | V |  | - |  |  |  | -5.4 |  | -5.4 |  |  | -5.4 |  |  |  | -5.4 |  |  | 0 |  |  | 0.0064 | -5.3936 | -5.4064 | PASS | CoC | 1.8e-05 | V | 2 | 9e-06 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | 1.0878e-15 | V | 1.7321 | 6.2804e-16 | 0.00028882 | 200.3888 | 2 | 0.00057763 | V | 1e-05 | V | 0.00057763 | V
|  | 60 V | [5 empty] | 6 |  | V |  | - |  |  |  | 6 |  | 6 |  |  | 6 |  |  |  | 6 |  |  | 0 |  |  | 0.016 | 6.016 | 5.984 | PASS | CoC | 1.8e-05 | V | 2 | 9e-06 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.0028868 | 200.0039 | 2 | 0.0057735 | V | 1e-05 | V | 0.0057735 | V
[8 empty] | 30 |  | V |  | - |  |  |  | 30 |  | 30 |  |  | 30 |  |  |  | 30 |  |  | 0 |  |  | 0.04 | 30.04 | 29.96 | PASS | CoC | 0.0001 | V | 2 | 5e-05 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.0028872 | 200.12 | 2 | 0.0057744 | V | 1e-05 | V | 0.0057744 | V
[8 empty] | 54 |  | V |  | - |  |  |  | 54 |  | 54 |  |  | 54 |  |  |  | 54 |  |  | 0 |  |  | 0.064 | 54.064 | 53.936 | PASS | CoC | 0.00031 | V | 2 | 0.000155 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.0028909 | 201.1532 | 2 | 0.0057818 | V | 0.00018 | V | 0.0057818 | V
[8 empty] | -6 |  | V |  | - |  |  |  | -6 |  | -6 |  |  | -6 |  |  |  | -6 |  |  | 0 |  |  | 0.016 | -5.984 | -6.016 | PASS | CoC | 1.8e-05 | V | 2 | 9e-06 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.0028868 | 200.0039 | 2 | 0.0057735 | V | 1e-05 | V | 0.0057735 | V
[8 empty] | -54 |  | V |  | - |  |  |  | -54 |  | -54 |  |  | -54 |  |  |  | -54 |  |  | 0 |  |  | 0.064 | -53.936 | -54.064 | PASS | CoC | 0.00031 | V | 2 | 0.000155 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.0028909 | 201.1532 | 2 | 0.0057818 | V | 0.00018 | V | 0.0057818 | V
|  | 600 V | [5 empty] | 60 |  | V |  | - |  |  |  | 60 |  | 60 |  |  | 60 |  |  |  | 60 |  |  | 0 |  |  | 0.16 | 60.16 | 59.84 | PASS | CoC | 0.00031 | V | 2 | 0.000155 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.028868 | 200.0115 | 2 | 0.057736 | V | 0.00018 | V | 0.057736 | V
[8 empty] | 540 |  | V |  | - |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 0.64 | 540.64 | 539.36 | PASS | CoC | 0.0033 | V | 2 | 0.00165 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.028915 | 201.3068 | 2 | 0.057829 | V | 0.0014 | V | 0.057829 | V
[8 empty] | -540 |  | V |  | - |  |  |  | -540 |  | -540 |  |  | -540 |  |  |  | -540 |  |  | 0 |  |  | 0.64 | -539.36 | -540.64 | PASS | CoC | 0.0033 | V | 2 | 0.00165 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.028915 | 201.3068 | 2 | 0.057829 | V | 0.0014 | V | 0.057829 | V
|  | 1000 V | [5 empty] | 100 |  | V |  | - |  |  |  | 100 |  | 100 |  |  | 100 |  |  |  | 100 |  |  | 0 |  |  | 1.1 | 101.1 | 98.9 | PASS | CoC | 0.00031 | V | 2 | 0.000155 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | 0 | V | 1.7321 | 0 | 0.28868 | 200.0001 | 2 | 0.57735 | V | 0.00018 | V | 0.57735 | V
[8 empty] | 900 |  | V |  | - |  |  |  | 900 |  | 900 |  |  | 900 |  |  |  | 900 |  |  | 0 |  |  | 1.9 | 901.9 | 898.1 | PASS | CoC | 0.0033 | V | 2 | 0.00165 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | 0 | V | 1.7321 | 0 | 0.28868 | 200.0131 | 2 | 0.57736 | V | 0.0014 | V | 0.57736 | V
[8 empty] | -900 |  | V |  | - |  |  |  | -900 |  | -900 |  |  | -900 |  |  |  | -900 |  |  | 0 |  |  | 1.9 | -898.1 | -901.9 | PASS | CoC | 0.0033 | V | 2 | 0.00165 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | 0 | V | 1.7321 | 0 | 0.28868 | 200.0131 | 2 | 0.57736 | V | 0.0014 | V | 0.57736 | V
|  | AC Voltage Test | [28 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
|  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  |  | Measured Value | [8 empty] | Average |  |  | Error |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  | 600 mV | [5 empty] | 60 |  | mV |  | 50 Hz |  |  |  | 60 |  | 60 |  |  | 60 |  |  |  | 60 |  |  | 0 |  |  | 0.82 | 60.82 | 59.18 | PASS | CoC | 0.07 | mV | 2 | 0.035 | Annual Drift | 0 | mV | 1.7321 | 0 | Readability | 0.1 | mV | 3.4641 | 0.028868 | Repeatability | 0 | mV | 1.7321 | 0 | 0.045369 | 386.023 | 2 | 0.090738 | mV | 0.01 | mV | 0.090738 | mV
[8 empty] | 60 |  | mV |  | 1 kHz |  |  |  | 60 |  | 60 |  |  | 60 |  |  |  | 60 |  |  | 0 |  |  | 1 | 61 | 59 | PASS | CoC | 0.07 | mV | 2 | 0.035 | Annual Drift | 0 | mV | 1.7321 | 0 | Readability | 0.1 | mV | 3.4641 | 0.028868 | Repeatability | 0 | mV | 1.7321 | 0 | 0.045369 | 386.023 | 2 | 0.090738 | mV | 0.01 | mV | 0.090738 | mV
[8 empty] | 540 |  | mV |  | 50 Hz |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 4.18 | 544.18 | 535.82 | PASS | CoC | 0.91 | mV | 2 | 0.455 | Annual Drift | 0 | mV | 1.7321 | 0 | Readability | 0.1 | mV | 3.4641 | 0.028868 | Repeatability | 0 | mV | 1.7321 | 0 | 0.45591 | 201.6101 | 2 | 0.91183 | mV | 0.085 | mV | 0.91183 | mV
[8 empty] | 540 |  | mV |  | 1 kHz |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 5.8 | 545.8 | 534.2 | PASS | CoC | 0.91 | mV | 2 | 0.455 | Annual Drift | 0 | mV | 1.7321 | 0 | Readability | 0.1 | mV | 3.4641 | 0.028868 | Repeatability | 0 | mV | 1.7321 | 0 | 0.45591 | 201.6101 | 2 | 0.91183 | mV | 0.085 | mV | 0.91183 | mV
|  | 6 V | [5 empty] | 0.6 |  | V |  | 50 Hz |  |  |  | 0.6 |  | 0.6 |  |  | 0.6 |  |  |  | 0.6 |  |  | 0 |  |  | 0.0062 | 0.6062 | 0.5938 | PASS | CoC | 0.00091 | V | 2 | 0.000455 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.00053885 | 338.5604 | 2 | 0.0010777 | V | 8.5e-05 | V | 0.0010777 | V
[8 empty] | 0.6 |  | V |  | 1 kHz |  |  |  | 0.6 |  | 0.6 |  |  | 0.6 |  |  |  | 0.6 |  |  | 0 |  |  | 0.01 | 0.61 | 0.59 | PASS | CoC | 0.00091 | V | 2 | 0.000455 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.00053885 | 338.5604 | 2 | 0.0010777 | V | 8.5e-05 | V | 0.0010777 | V
[8 empty] | 5.4 |  | V |  | 50 Hz |  |  |  | 5.4 |  | 5.4 |  |  | 5.4 |  |  |  | 5.4 |  |  | 0 |  |  | 0.0398 | 5.4398 | 5.3602 | PASS | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | 1.0878e-15 | V | 1.7321 | 6.2804e-16 | 0.0036614 | 202.5019 | 2 | 0.0073228 | V | 0.00085 | V | 0.0073228 | V
[8 empty] | 5.4 |  | V |  | 1 kHz |  |  |  | 5.4 |  | 5.4 |  |  | 5.4 |  |  |  | 5.4 |  |  | 0 |  |  | 0.058 | 5.458 | 5.342 | PASS | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | 1.0878e-15 | V | 1.7321 | 6.2804e-16 | 0.0036614 | 202.5019 | 2 | 0.0073228 | V | 0.00085 | V | 0.0073228 | V
|  | 60 V | [5 empty] | 6 |  | V |  | 50 Hz |  |  |  | 6 |  | 6 |  |  | 6 |  |  |  | 6 |  |  | 0 |  |  | 0.062 | 6.062 | 5.938 | PASS | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.0046536 | 379.8393 | 2 | 0.0093072 | V | 0.00085 | V | 0.0093072 | V
[8 empty] | 6 |  | V |  | 1 kHz |  |  |  | 6 |  | 6 |  |  | 6 |  |  |  | 6 |  |  | 0 |  |  | 0.1 | 6.1 | 5.9 | PASS | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.0046536 | 379.8393 | 2 | 0.0093072 | V | 0.00085 | V | 0.0093072 | V
[8 empty] | 30 |  | V |  | 50 Hz |  |  |  | 30 |  | 30 |  |  | 30 |  |  |  | 30 |  |  | 0 |  |  | 0.23 | 30.23 | 29.77 | PASS | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.0046536 | 379.8393 | 2 | 0.0093072 | V | 0.00085 | V | 0.0093072 | V
[8 empty] | 30 |  | V |  | 1 kHz |  |  |  | 30 |  | 30 |  |  | 30 |  |  |  | 30 |  |  | 0 |  |  | 0.34 | 30.34 | 29.66 | PASS | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.0046536 | 379.8393 | 2 | 0.0093072 | V | 0.00085 | V | 0.0093072 | V
[8 empty] | 54 |  | V |  | 50 Hz |  |  |  | 54 |  | 54 |  |  | 54 |  |  |  | 54 |  |  | 0 |  |  | 0.398 | 54.398 | 53.602 | PASS | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.035617 | 202.6449 | 2 | 0.071234 | V | 0.0064 | V | 0.071234 | V
[8 empty] | 54 |  | V |  | 1 kHz |  |  |  | 54 |  | 54 |  |  | 54 |  |  |  | 54 |  |  | 0 |  |  | 0.58 | 54.58 | 53.42 | PASS | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.035617 | 202.6449 | 2 | 0.071234 | V | 0.0064 | V | 0.071234 | V
|  | 600 V | [5 empty] | 60 |  | V |  | 50 Hz |  |  |  | 60 |  | 60 |  |  | 60 |  |  |  | 60 |  |  | 0 |  |  | 0.62 | 60.62 | 59.38 | PASS | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.045756 | 384.0312 | 2 | 0.091511 | V | 0.0064 | V | 0.091511 | V
[8 empty] | 60 |  | V |  | 1 kHz |  |  |  | 60 |  | 60 |  |  | 60 |  |  |  | 60 |  |  | 0 |  |  | 1 | 61 | 59 | PASS | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.045756 | 384.0312 | 2 | 0.091511 | V | 0.0064 | V | 0.091511 | V
[8 empty] | 540 |  | V |  | 50 Hz |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 3.98 | 543.98 | 536.02 | PASS | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.045756 | 384.0312 | 2 | 0.091511 | V | 0.084 | V | 0.091511 | V
[8 empty] | 540 |  | V |  | 1 kHz |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 5.8 | 545.8 | 534.2 | PASS | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | 0 | V | 1.7321 | 0 | 0.045756 | 384.0312 | 2 | 0.091511 | V | 0.084 | V | 0.091511 | V
|  | 1000 V | [5 empty] | 100 |  | V |  | 50 Hz |  |  |  | 100 |  | 100 |  |  | 100 |  |  |  | 100 |  |  | 0 |  |  | 2.7 | 102.7 | 97.3 | PASS | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | 0 | V | 1.7321 | 0 | 0.29085 | 206.0478 | 2 | 0.5817 | V | 0.0064 | V | 0.5817 | V
[8 empty] | 100 |  | V |  | 1 kHz |  |  |  | 100 |  | 100 |  |  | 100 |  |  |  | 100 |  |  | 0 |  |  | 5 | 105 | 95 | PASS | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | 0 | V | 1.7321 | 0 | 0.29085 | 206.0478 | 2 | 0.5817 | V | 0.0064 | V | 0.5817 | V
[8 empty] | 900 |  | V |  | 50 Hz |  |  |  | 900 |  | 900 |  |  | 900 |  |  |  | 900 |  |  | 0 |  |  | 8.3 | 908.3 | 891.7 | PASS | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | 0 | V | 1.7321 | 0 | 0.29085 | 206.0478 | 2 | 0.5817 | V | 0.084 | V | 0.5817 | V
[8 empty] | 900 |  | V |  | 1 kHz |  |  |  | 900 |  | 900 |  |  | 900 |  |  |  | 900 |  |  | 0 |  |  | 13 | 913 | 887 | PASS | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | 0 | V | 1.7321 | 0 | 0.29085 | 206.0478 | 2 | 0.5817 | V | 0.084 | V | 0.5817 | V
|  | Resistance Test | [28 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
|  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  |  | Measured Value | [8 empty] | Average |  |  | Error |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  | 600 Ω | [5 empty] | 0 |  | Ω |  | - |  |  |  | 0 |  | 0 |  |  | 0 |  |  |  | 0 |  |  | 0 |  |  | 0.2 | 0.2 | -0.2 | PASS | CoC | 7.7e-05 | Ω | 2 | 3.85e-05 | Annual Drift | 0 | Ω | 1.7321 | 0 | Readability | 0.1 | Ω | 3.4641 | 0.028868 | Repeatability | 0 | Ω | 1.7321 | 0 | 0.028868 | 200.0007 | 2 | 0.057735 | Ω | 0.0078 | Ω | 0.057735 | Ω
[8 empty] | 540 |  | Ω |  | - |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 1.28 | 541.28 | 538.72 | PASS | CoC | 0.003 | Ω | 2 | 0.0015 | Annual Drift | 0 | Ω | 1.7321 | 0 | Readability | 0.1 | Ω | 3.4641 | 0.028868 | Repeatability | 0 | Ω | 1.7321 | 0 | 0.028906 | 201.08 | 2 | 0.057813 | Ω | 0.023 | Ω | 0.057813 | Ω
|  | 6 kΩ | [5 empty] | 5.4 |  | kΩ |  | - |  |  |  | 5.4 |  | 5.4 |  |  | 5.4 |  |  |  | 5.4 |  |  | 0 |  |  | 0.0118 | 5.4118 | 5.3882 | PASS | CoC | 5.5e-05 | kΩ | 2 | 2.75e-05 | Annual Drift | 0 | kΩ | 1.7321 | 0 | Readability | 0.001 | kΩ | 3.4641 | 0.00028868 | Repeatability | 1.0878e-15 | kΩ | 1.7321 | 6.2804e-16 | 0.00028998 | 203.6297 | 2 | 0.00057996 | kΩ | 0.00023 | kΩ | 0.00057996 | kΩ
|  | 60 kΩ | [5 empty] | 54 |  | kΩ |  | - |  |  |  | 54 |  | 54 |  |  | 54 |  |  |  | 54 |  |  | 0 |  |  | 0.118 | 54.118 | 53.882 | PASS | CoC | 0.00081 | kΩ | 2 | 0.000405 | Annual Drift | 0 | kΩ | 1.7321 | 0 | Readability | 0.01 | kΩ | 3.4641 | 0.0028868 | Repeatability | 0 | kΩ | 1.7321 | 0 | 0.002915 | 207.8702 | 2 | 0.00583 | kΩ | 0.0015 | kΩ | 0.00583 | kΩ
|  | 600 kΩ | [5 empty] | 540 |  | kΩ |  | - |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 3.34 | 543.34 | 536.66 | PASS | CoC | 0.0066 | kΩ | 2 | 0.0033 | Annual Drift | 0 | kΩ | 1.7321 | 0 | Readability | 0.1 | kΩ | 3.4641 | 0.028868 | Repeatability | 0 | kΩ | 1.7321 | 0 | 0.029056 | 205.2263 | 2 | 0.058111 | kΩ | 0.016 | kΩ | 0.058111 | kΩ
|  | 6 MΩ | [5 empty] | 5.4 |  | MΩ |  | - |  |  |  | 5.4 |  | 5.4 |  |  | 5.4 |  |  |  | 5.4 |  |  | 0 |  |  | 0.0334 | 5.4334 | 5.3666 | PASS | CoC | 0.00048 | MΩ | 2 | 0.00024 | Annual Drift | 0 | MΩ | 1.7321 | 0 | Readability | 0.001 | MΩ | 3.4641 | 0.00028868 | Repeatability | 1.0878e-15 | MΩ | 1.7321 | 6.2804e-16 | 0.00037541 | 387.0943 | 2 | 0.00075082 | MΩ | 0.00053 | MΩ | 0.00075082 | MΩ
|  | 50 MΩ | [5 empty] | 45 |  | MΩ |  | - |  |  |  | 45 |  | 45 |  |  | 45 |  |  |  | 45 |  |  | 0 |  |  | 0.55 | 45.55 | 44.45 | PASS | CoC | 0.0047 | MΩ | 2 | 0.00235 | Annual Drift | 0 | MΩ | 1.7321 | 0 | Readability | 0.01 | MΩ | 3.4641 | 0.0028868 | Repeatability | 0 | MΩ | 1.7321 | 0 | 0.0037223 | 384.1893 | 2 | 0.0074447 | MΩ | 0.015 | MΩ | 0.015 | MΩ
|  | DC Current Test | [28 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
|  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  |  | Measured Value | [8 empty] | Average |  |  | Error |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  | 600 µA | [5 empty] | 540 |  | µA |  | - |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 1.48 | 541.48 | 538.52 | PASS | CoC | 0.09 | µA | 2 | 0.045 | Annual Drift | 0 | µA | 1.7321 | 0 | Readability | 0.1 | µA | 3.4641 | 0.028868 | Repeatability | 0 | µA | 1.7321 | 0 | 0.053463 | 340.7696 | 2 | 0.10693 | µA | 0.064 | µA | 0.10693 | µA
|  | 6000 µA | [5 empty] | 5400 |  | µA |  | - |  |  |  | 5400 |  | 5400 |  |  | 5400 |  |  |  | 5400 |  |  | 0 |  |  | 12.8 | 5412.8 | 5387.2 | PASS | CoC | 0.38 | µA | 2 | 0.19 | Annual Drift | 0 | µA | 1.7321 | 0 | Readability | 1 | µA | 3.4641 | 0.28868 | Repeatability | 0 | µA | 1.7321 | 0 | 0.34559 | 345.9001 | 2 | 0.69118 | µA | 0.45 | µA | 0.69118 | µA
|  | 60 mA | [5 empty] | 54 |  | mA |  | - |  |  |  | 54 |  | 54 |  |  | 54 |  |  |  | 54 |  |  | 0 |  |  | 0.148 | 54.148 | 53.852 | PASS | CoC | 0.0014 | mA | 2 | 0.0007 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Repeatability | 0 | mA | 1.7321 | 0 | 0.0029704 | 223.439 | 2 | 0.0059408 | mA | 0.0043 | mA | 0.0059408 | mA
[8 empty] | -54 |  | mA |  | - |  |  |  | -54 |  | -54 |  |  | -54 |  |  |  | -54 |  |  | 0 |  |  | 0.148 | -53.852 | -54.148 | PASS | CoC | 0.0014 | mA | 2 | 0.0007 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Repeatability | 0 | mA | 1.7321 | 0 | 0.0029704 | 223.439 | 2 | 0.0059408 | mA | 0.0043 | mA | 0.0059408 | mA
|  | 400 mA | [5 empty] | 360 |  | mA |  | - |  |  |  | 360 |  | 360 |  |  | 360 |  |  |  | 360 |  |  | 0 |  |  | 0.92 | 360.92 | 359.08 | PASS | CoC | 0.027 | mA | 2 | 0.0135 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.1 | mA | 3.4641 | 0.028868 | Repeatability | 0 | mA | 1.7321 | 0 | 0.031868 | 283.4868 | 2 | 0.063736 | mA | 0.082 | mA | 0.082 | mA
|  | 6 A | [5 empty] | 3 |  | A |  | - |  |  |  | 3 |  | 3 |  |  | 3 |  |  |  | 3 |  |  | 0 |  |  | 0.01 | 3.01 | 2.99 | PASS | CoC | 0.0014 | A | 2 | 0.0007 | Annual Drift | 0 | A | 1.7321 | 0 | Readability | 0.001 | A | 3.4641 | 0.00028868 | Repeatability | 0 | A | 1.7321 | 0 | 0.00075719 | 266.115 | 2 | 0.0015144 | A | 0.0016 | A | 0.0016 | A
[8 empty] | 5.4 |  | A |  | - |  |  |  | 5.4 |  | 5.4 |  |  | 5.4 |  |  |  | 5.4 |  |  | 0 |  |  | 0.0148 | 5.4148 | 5.3852 | PASS | CoC | 0.0014 | A | 2 | 0.0007 | Annual Drift | 0 | A | 1.7321 | 0 | Readability | 0.001 | A | 3.4641 | 0.00028868 | Repeatability | 1.0878e-15 | A | 1.7321 | 6.2804e-16 | 0.00075719 | 266.115 | 2 | 0.0015144 | A | 0.0016 | A | 0.0016 | A
|  | 10 A | [5 empty] | 5 |  | A |  | - |  |  |  | 5 |  | 5 |  |  | 5 |  |  |  | 5 |  |  | 0 |  |  | 0.03 | 5.03 | 4.97 | PASS | CoC | 0.0014 | A | 2 | 0.0007 | Annual Drift | 0 | A | 1.7321 | 0 | Readability | 0.01 | A | 3.4641 | 0.0028868 | Repeatability | 0 | A | 1.7321 | 0 | 0.0029704 | 223.439 | 2 | 0.0059408 | A | 0.0016 | A | 0.0059408 | A
[8 empty] | 9 |  | A |  | - |  |  |  | 9 |  | 9 |  |  | 9 |  |  |  | 9 |  |  | 0 |  |  | 0.038 | 9.038 | 8.962 | PASS | CoC | 0.0014 | A | 2 | 0.0007 | Annual Drift | 0 | A | 1.7321 | 0 | Readability | 0.01 | A | 3.4641 | 0.0028868 | Repeatability | 0 | A | 1.7321 | 0 | 0.0029704 | 223.439 | 2 | 0.0059408 | A | 0.0016 | A | 0.0059408 | A
|  | AC Current Test | [28 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
|  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  |  | Measured Value | [8 empty] | Average |  |  | Error |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  | 600 µA | [5 empty] | 540 |  | µA |  | 50 Hz |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 5.6 | 545.6 | 534.4 | PASS | CoC | 1.7 | µA | 2 | 0.85 | Annual Drift | 0 | µA | 1.7321 | 0 | Readability | 0.1 | µA | 3.4641 | 0.028868 | Repeatability | 0 | µA | 1.7321 | 0 | 0.85049 | 200.4614 | 2 | 1.701 | µA | 0.00037 | µA | 1.701 | µA
[8 empty] | 540 |  | µA |  | 1 kHz |  |  |  | 540 |  | 540 |  |  | 540 |  |  |  | 540 |  |  | 0 |  |  | 5.6 | 545.6 | 534.4 | PASS | CoC | 1.7 | µA | 2 | 0.85 | Annual Drift | 0 | µA | 1.7321 | 0 | Readability | 0.1 | µA | 3.4641 | 0.028868 | Repeatability | 0 | µA | 1.7321 | 0 | 0.85049 | 200.4614 | 2 | 1.701 | µA | 0.00037 | µA | 1.701 | µA
|  | 6000 µA | [5 empty] | 5400 |  | µA |  | 50 Hz |  |  |  | 5400 |  | 5400 |  |  | 5400 |  |  |  | 5400 |  |  | 0 |  |  | 56 | 5456 | 5344 | PASS | CoC | 13 | µA | 2 | 6.5 | Annual Drift | 0 | µA | 1.7321 | 0 | Readability | 1 | µA | 3.4641 | 0.28868 | Repeatability | 0 | µA | 1.7321 | 0 | 6.5064 | 200.789 | 2 | 13.0128 | µA | 2.6 | µA | 13.0128 | µA
[8 empty] | 5400 |  | µA |  | 1 kHz |  |  |  | 5400 |  | 5400 |  |  | 5400 |  |  |  | 5400 |  |  | 0 |  |  | 56 | 5456 | 5344 | PASS | CoC | 13 | µA | 2 | 6.5 | Annual Drift | 0 | µA | 1.7321 | 0 | Readability | 1 | µA | 3.4641 | 0.28868 | Repeatability | 0 | µA | 1.7321 | 0 | 6.5064 | 200.789 | 2 | 13.0128 | µA | 2.6 | µA | 13.0128 | µA
|  | 60 mA | [5 empty] | 54 |  | mA |  | 50 Hz |  |  |  | 54 |  | 54 |  |  | 54 |  |  |  | 54 |  |  | 0 |  |  | 0.56 | 54.56 | 53.44 | PASS | CoC | 0.22 | mA | 2 | 0.11 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Repeatability | 0 | mA | 1.7321 | 0 | 0.11004 | 200.2755 | 2 | 0.22008 | mA | 0.026 | mA | 0.22008 | mA
[8 empty] | 54 |  | mA |  | 1 kHz |  |  |  | 54 |  | 54 |  |  | 54 |  |  |  | 54 |  |  | 0 |  |  | 0.56 | 54.56 | 53.44 | PASS | CoC | 0.22 | mA | 2 | 0.11 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Repeatability | 0 | mA | 1.7321 | 0 | 0.11004 | 200.2755 | 2 | 0.22008 | mA | 0.026 | mA | 0.22008 | mA
|  | 400 mA | [5 empty] | 360 |  | mA |  | 50 Hz |  |  |  | 360 |  | 360 |  |  | 360 |  |  |  | 360 |  |  | 0 |  |  | 3.8 | 363.8 | 356.2 | PASS | CoC | 0.72 | mA | 2 | 0.36 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.1 | mA | 3.4641 | 0.028868 | Repeatability | 0 | mA | 1.7321 | 0 | 0.36116 | 202.5719 | 2 | 0.72231 | mA | 0.21 | mA | 0.72231 | mA
[8 empty] | 360 |  | mA |  | 1 kHz |  |  |  | 360 |  | 360 |  |  | 360 |  |  |  | 360 |  |  | 0 |  |  | 3.8 | 363.8 | 356.2 | PASS | CoC | 0.72 | mA | 2 | 0.36 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.1 | mA | 3.4641 | 0.028868 | Repeatability | 0 | mA | 1.7321 | 0 | 0.36116 | 202.5719 | 2 | 0.72231 | mA | 0.21 | mA | 0.72231 | mA
|  | 6 A | [5 empty] | 5.4 |  | A |  | 50 Hz |  |  |  | 5.4 |  | 5.4 |  |  | 5.4 |  |  |  | 5.4 |  |  | 0 |  |  | 0.056 | 5.456 | 5.344 | PASS | CoC | 0.01 | A | 2 | 0.005 | Annual Drift | 0 | A | 1.7321 | 0 | Readability | 0.001 | A | 3.4641 | 0.00028868 | Repeatability | 1.0878e-15 | A | 1.7321 | 6.2804e-16 | 0.0050083 | 201.3333 | 2 | 0.010017 | A | 0.0029 | A | 0.010017 | A
[8 empty] | 5.4 |  | A |  | 1 kHz |  |  |  | 5.4 |  | 5.4 |  |  | 5.4 |  |  |  | 5.4 |  |  | 0 |  |  | 0.056 | 5.456 | 5.344 | PASS | CoC | 0.011 | A | 2 | 0.0055 | Annual Drift | 0 | A | 1.7321 | 0 | Readability | 0.001 | A | 3.4641 | 0.00028868 | Repeatability | 1.0878e-15 | A | 1.7321 | 6.2804e-16 | 0.0055076 | 201.1019 | 2 | 0.011015 | A | 0.0039 | A | 0.011015 | A
|  | 10 A | [5 empty] | 9 |  | A |  | 50 Hz |  |  |  | 9 |  | 9 |  |  | 9 |  |  |  | 9 |  |  | 0 |  |  | 0.11 | 9.11 | 8.89 | PASS | CoC | 0.01 | A | 2 | 0.005 | Annual Drift | 0 | A | 1.7321 | 0 | Readability | 0.01 | A | 3.4641 | 0.0028868 | Repeatability | 0 | A | 1.7321 | 0 | 0.0057735 | 320 | 2 | 0.011547 | A | 0.0029 | A | 0.011547 | A
[8 empty] | 9 |  | A |  | 1 kHz |  |  |  | 9 |  | 9 |  |  | 9 |  |  |  | 9 |  |  | 0 |  |  | 0.11 | 9.11 | 8.89 | PASS | CoC | 0.011 | A | 2 | 0.0055 | Annual Drift | 0 | A | 1.7321 | 0 | Readability | 0.01 | A | 3.4641 | 0.0028868 | Repeatability | 0 | A | 1.7321 | 0 | 0.0062115 | 302.4201 | 2 | 0.012423 | A | 0.0039 | A | 0.012423 | A
Calibrated By
DANICA BEA NERI | [22 empty] | 2025-04-09
Calibration Officer's Name | [13 empty] | Signature | [8 empty] | Date
Checked By
Jennefer S. Ilas
Interim Manager (Electrical) | [13 empty] | Signature | [8 empty] | Date
Encoded By
- | [13 empty] | - | [8 empty] | -
Calibration Officer's Name | [13 empty] | Signature | [8 empty] | Date

## PASSFAIL (hidden)
PASS/FAIL ANALYSIS
Measurement Results | [5 empty] | Based on Tolerance |  |  |  |  | Based on +/- Uncertainty Applied | [7 empty] | FINAL REMARKS
DC Voltage Test
Measured Value |  | Tolerance |  | Uncertainty |  | Upper Limit |  | Lower Limit |  | Remarks | Measured Value + Uncertainty |  | Upper Limit |  | Remarks | Measured Value - Uncertainty |  | Remarks
60 | mV | 0.72 | mV | 0.05774 | mV | 60.72 | mV | 59.28 | mV | PASSED | 60.0577 | mV | 60.72 | mV | PASSED | 59.9423 | mV | PASSED | PASSED
540 | mV | 4.08 | mV | 0.057773 | mV | 544.08 | mV | 535.92 | mV | PASSED | 540.0578 | mV | 544.08 | mV | PASSED | 539.9422 | mV | PASSED | PASSED
-540 | mV | 4.08 | mV | 0.057773 | mV | -535.92 | mV | -544.08 | mV | PASSED | -539.9422 | mV | -535.92 | mV | PASSED | -540.0578 | mV | PASSED | PASSED
0.6 | V | 0.0072 | V | 0.00057735 | V | 0.6072 | V | 0.5928 | V | PASSED | 0.60058 | V | 0.6072 | V | PASSED | 0.59942 | V | PASSED | PASSED
5.4 | V | 0.0408 | V | 0.00057763 | V | 5.4408 | V | 5.3592 | V | PASSED | 5.4006 | V | 5.4408 | V | PASSED | 5.3994 | V | PASSED | PASSED
-5.4 | V | 0.0408 | V | 0.00057763 | V | -5.3592 | V | -5.4408 | V | PASSED | -5.3994 | V | -5.3592 | V | PASSED | -5.4006 | V | PASSED | PASSED
6 | V | 0.072 | V | 0.0057735 | V | 6.072 | V | 5.928 | V | PASSED | 6.0058 | V | 6.072 | V | PASSED | 5.9942 | V | PASSED | PASSED
30 | V | 0.24 | V | 0.0057744 | V | 30.24 | V | 29.76 | V | PASSED | 30.0058 | V | 30.24 | V | PASSED | 29.9942 | V | PASSED | PASSED
54 | V | 0.408 | V | 0.0057818 | V | 54.408 | V | 53.592 | V | PASSED | 54.0058 | V | 54.408 | V | PASSED | 53.9942 | V | PASSED | PASSED
-6 | V | 0.072 | V | 0.0057735 | V | -5.928 | V | -6.072 | V | PASSED | -5.9942 | V | -5.928 | V | PASSED | -6.0058 | V | PASSED | PASSED
-54 | V | 0.408 | V | 0.0057818 | V | -53.592 | V | -54.408 | V | PASSED | -53.9942 | V | -53.592 | V | PASSED | -54.0058 | V | PASSED | PASSED
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
60 | V | 0.72 | V | 0.057736 | V | 60.72 | V | 59.28 | V | PASSED | 60.0577 | V | 60.72 | V | PASSED | 59.9423 | V | PASSED | PASSED
540 | V | 4.08 | V | 0.057829 | V | 544.08 | V | 535.92 | V | PASSED | 540.0578 | V | 544.08 | V | PASSED | 539.9422 | V | PASSED | PASSED
-540 | V | 4.08 | V | 0.057829 | V | -535.92 | V | -544.08 | V | PASSED | -539.9422 | V | -535.92 | V | PASSED | -540.0578 | V | PASSED | PASSED
AC Voltage Test | [18 empty] | FINAL REMARKS
Measured Value |  | Tolerance |  | Uncertainty |  | Upper Limit |  | Lower Limit |  | Remarks | Measured + Upper UNC |  | Upper Limit |  | Remarks | Measured + Lower UNC |  | Remarks
0.6 | V | 0.0186 | V | 0.0010777 | V | 0.6186 |  | 0.5814 |  | PASSED | 0.60108 | V | 0.6186 | V | PASSED | 0.59892 | V | PASSED | PASSED
0.6 | V | 0.0186 | V | 0.0010777 | V | 0.6186 |  | 0.5814 |  | PASSED | 0.60108 | V | 0.6186 | V | PASSED | 0.59892 | V | PASSED | PASSED
5.4 | V | 0.0954 | V | 0.0073228 | V | 5.4954 |  | 5.3046 |  | PASSED | 5.4073 | V | 5.4954 | V | PASSED | 5.3927 | V | PASSED | PASSED
5.4 | V | 0.0954 | V | 0.0073228 | V | 5.4954 |  | 5.3046 |  | PASSED | 5.4073 | V | 5.4954 | V | PASSED | 5.3927 | V | PASSED | PASSED
6 | V | 0.186 | V | 0.0093072 | V | 6.186 |  | 5.814 |  | PASSED | 6.0093 | V | 6.186 | V | PASSED | 5.9907 | V | PASSED | PASSED
6 | V | 0.186 | V | 0.0093072 | V | 6.186 |  | 5.814 |  | PASSED | 6.0093 | V | 6.186 | V | PASSED | 5.9907 | V | PASSED | PASSED
30 | V | 0.57 | V | 0.0093072 | V | 30.57 |  | 29.43 |  | PASSED | 30.0093 | V | 30.57 | V | PASSED | 29.9907 | V | PASSED | PASSED
30 | V | 0.57 | V | 0.0093072 | V | 30.57 |  | 29.43 |  | PASSED | 30.0093 | V | 30.57 | V | PASSED | 29.9907 | V | PASSED | PASSED
54 | V | 0.954 | V | 0.071234 | V | 54.954 |  | 53.046 |  | PASSED | 54.0712 | V | 54.954 | V | PASSED | 53.9288 | V | PASSED | PASSED
54 | V | 0.954 | V | 0.071234 | V | 54.954 |  | 53.046 |  | PASSED | 54.0712 | V | 54.954 | V | PASSED | 53.9288 | V | PASSED | PASSED
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | #REF! |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | #REF! |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | #REF! |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | #REF! |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
60 | V | 1.86 | V | 0.091511 | V | 61.86 |  | 58.14 |  | PASSED | 60.0915 | V | 61.86 | V | PASSED | 59.9085 | V | PASSED | PASSED
60 | V | 1.86 | V | 0.091511 | V | 61.86 |  | 58.14 |  | PASSED | 60.0915 | V | 61.86 | V | PASSED | 59.9085 | V | PASSED | PASSED
540 | V | 9.54 | V | 0.091511 | V | 549.54 |  | 530.46 |  | PASSED | 540.0915 | V | 549.54 | V | PASSED | 539.9085 | V | PASSED | PASSED
540 | V | 9.54 | V | 0.091511 | V | 549.54 |  | 530.46 |  | PASSED | 540.0915 | V | 549.54 | V | PASSED | 539.9085 | V | PASSED | PASSED
Resistance Test | [18 empty] | FINAL REMARKS
Measured Value |  | Tolerance |  | Uncertainty |  | Upper Limit |  | Lower Limit |  | Remarks | Measured + Upper UNC |  | Upper Limit |  | Remarks | Measured + Lower UNC |  | Remarks
0 | Ω | 0.5 | Ω | 0.057735 | Ω | 0.5 |  | -0.5 |  | PASSED | 0.057735 | Ω | 0.5 | Ω | PASSED | -0.057735 | Ω | PASSED | PASSED
540 | Ω | 8.6 | Ω | 0.057813 | Ω | 548.6 |  | 531.4 |  | PASSED | 540.0578 | Ω | 548.6 | Ω | PASSED | 539.9422 | Ω | PASSED | PASSED
5.4 | kΩ | 0.086 | kΩ | 0.00057996 | kΩ | 5.486 |  | 5.314 |  | PASSED | 5.4006 | kΩ | 5.486 | kΩ | PASSED | 5.3994 | kΩ | PASSED | PASSED
54 | kΩ | 0.86 | kΩ | 0.00583 | kΩ | 54.86 |  | 53.14 |  | PASSED | 54.0058 | kΩ | 54.86 | kΩ | PASSED | 53.9942 | kΩ | PASSED | PASSED
540 | kΩ | 8.6 | kΩ | 0.058111 | kΩ | 548.6 |  | 531.4 |  | PASSED | 540.0581 | kΩ | 548.6 | kΩ | PASSED | 539.9419 | kΩ | PASSED | PASSED
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | #REF! |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
45 | MΩ | 0.725 | MΩ | 0.015 | MΩ | 45.725 |  | 44.275 |  | PASSED | 45.015 | MΩ | 45.725 | MΩ | PASSED | 44.985 | MΩ | PASSED | PASSED
DC Current Test | [18 empty] | FINAL REMARKS
Measured Value |  | Tolerance |  | Uncertainty |  | Upper Limit |  | Lower Limit |  | Remarks | Measured + Upper UNC |  | Upper Limit |  | Remarks | Measured + Lower UNC |  | Remarks
3 | A | 0.071 | A | 0.0016 | A | 3.071 |  | 2.929 |  | PASSED | 3.0016 | A | 3.071 | A | PASSED | 2.9984 | A | PASSED | PASSED
5.4 | A | 0.1138 | A | 0.0016 | A | 5.5138 |  | 5.2862 |  | PASSED | 5.4016 | A | 5.5138 | A | PASSED | 5.3984 | A | PASSED | PASSED
9 | A | 0.248 | A | 0.0059408 | A | 9.248 |  | 8.752 |  | PASSED | 9.0059 | A | 9.248 | A | PASSED | 8.9941 | A | PASSED | PASSED
AC Current Test | [18 empty] | FINAL REMARKS
Measured Value |  | Tolerance |  | Uncertainty |  | Upper Limit |  | Lower Limit |  | Remarks | Measured + Upper UNC |  | Upper Limit |  | Remarks | Measured + Lower UNC |  | Remarks
5.4 | A | 0.1562 | A | 0.010017 | A | 5.5562 |  | 5.2438 |  | PASSED | 5.41 | A | 5.5562 | A | PASSED | 5.39 | A | PASSED | PASSED
5.4 | A | 0.1562 | A | 0.011015 | A | 5.5562 |  | 5.2438 |  | PASSED | 5.411 | A | 5.5562 | A | PASSED | 5.389 | A | PASSED | PASSED
9 | A | 0.302 | A | 0.011547 | A | 9.302 |  | 8.698 |  | PASSED | 9.0115 | A | 9.302 | A | PASSED | 8.9885 | A | PASSED | PASSED
9 | A | 0.302 | A | 0.012423 | A | 9.302 |  | 8.698 |  | PASSED | 9.0124 | A | 9.302 | A | PASSED | 8.9876 | A | PASSED | PASSED