## Cal Certificate
Calibration Certificate
[7 empty] | 0021-08-08 |  |  |  |  | Certificate No.: |  |  |  |  | DBN
05/08/2025 | Customer : |  |  |  |  | DIGITAL MULTIMETER |  |  | Location of Calibration :
05/08/2025 | [5 empty] | FLUKE
NA
NA
NA
See Specification Sheet | [5 empty] | See Specification Sheet
DANICA BEA NERI | Manufacturer: | [7 empty] | Description:
| Model No: | [7 empty] | Options Installed:
| Serial No: | [7 empty] | Customer Asset No:
| Calibration Date: | [7 empty] | Work Order No.:
| Received Date: | [7 empty] | Customer P.O.:
| Temperature: | [7 empty] | Humidity:
Cal Method / Procedure :
This is to certify that the above Measuring and Test Equipment has been calibrated by SONJU Engineering Services in accordance with relevant calibration procedure defined for this instrument.
| Uncertainty: | [14 empty] | b
The expanded uncertainty of measurement for the results given in this certificate was estimated with a coverage factor of k = 2 at 95% level of confidence. The uncertainties of measurement were calculated using a budget worksheet based on JCGM 100:2008 Evaluation of measurement data - Guide to the Expression of Uncertainty in Measurement, 1st edition- 2008.
Traceability:
The instrument was calibrated using reference standard/s traceable to SI Units as maintained by NML-ITDI (Philippines) through SONJU Engineering Services (Philippines)
Reference Standard Used:
|  | Description |  |  | Serial No. |  |  |  |  | Cal Report Ref. |  |  |  |  | Due Date
Remarks:
1. The results were the average of the three (3) readings and were obtained during the time of calibration and refer only to the particular instrument submitted. Should the instrument be modified or damaged in any way, or develop inconsistent readings, the results may not be valid and the unit will require recalibration.
2. The user should determine the suitability of this instrument for its intended use.
3. Preliminary Inspection: No response at high and low current.
4. A positve error is to subtracted to the UUT indication and negative error added to give true value.
5. As per customer request the next calibration due date is on:
|  |  | Calibration Officer | [9 empty] | Approved Signatory
Issue Date :
Measurement Results:
|  |  | Range | Nominal Value |  |  |  | Measured Value |  |  | Error |  |  | Uncertainty
|  |  | Range | Nominal Value |  |  |  | Measured Value |  |  | Error |  |  | Uncertainty
Calibration Officer
Measurement Results:
|  |  | Range | Nominal Value |  |  |  | Measured Value |  |  | Error |  |  | Uncertainty
|  |  | Range | Nominal Value |  |  |  | Measured Value |  |  | Error |  |  | Uncertainty
Calibration Officer
Measurement Results:
|  |  | Range | Nominal Value |  |  |  | Measured Value |  |  | Error |  |  | Uncertainty
Calibration Officer
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049) 513-5233/591-0807;
(02) 782-2153; 09236581107
Email: admin@sonjuengineering.com
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049) 513-5233/591-0807;
(02) 782-2153; 09236581107
Email: admin@sonjuengineering.com
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049) 513-5233/591-0807;
(02) 782-2153; 09236581107
Email: admin@sonjuengineering.com
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049) 513-5233/591-0807;
(02) 782-2153; 09236581107
Email: admin@sonjuengineering.com
Website: www.sonjuengineering.com

## DataSheet
CALIBRATION DATASHEET
|  | Work Order No.: | [15 empty] | Technical ID:
Unit Under Calibration
Description |  |  |  |  | : | [8 empty] | Received Date |  |  |  |  | :
Manufacturer |  |  |  |  | : | [8 empty] | Calibration Date |  |  |  |  | :
Model |  |  |  |  | : | [8 empty] | Options Installed |  |  |  |  | :
Serial Number |  |  |  |  | : | [8 empty] | Customer P.O. |  |  |  |  | :
Range |  |  |  |  | : | [8 empty] | Asset Number |  |  |  |  | :
Res/Readability |  |  |  |  | : | [8 empty] | Accuracy |  |  |  |  | :
Prev. SES Cal Cert: |  |  |  |  | : | [8 empty] | Previous Technician |  |  |  |  | :
Customer Details | [13 empty] | Calibration Location
Company |  |  | : | [12 empty] | In-House (SONJU Laboratory)
Address |  |  | : | [12 empty] | On-Site Calibration (Please Specify)
Electrical Calibration Area
Reference Standard/s Used
Description | [8 empty] | Serial No. |  |  |  |  | Cal Report Ref. | [6 empty] | Due Date
Accessory Used
Description | [8 empty] | Serial No. |  |  |  |  | Brand | [6 empty] | Model
Environmental Condition | [13 empty] | Calibration Method
|  |  |  | Temperature |  |  |  |  | Relative Humidity |  |  |  |  | As per SONJU Engineering Services Work Instruction 7W-02-54 Digital Multimeter Calibration Procedure.
[5 empty] | Start: |  | oC |  |  | Start: |  | %rh
[5 empty] | End: |  | oC |  |  | End: |  | %rh
Observations
| Busted fused both high and low current | [21 empty] | Adjustment
| - | [21 empty] | Needed?
| - | [22 empty] | YES
| - | [22 empty] | NO
Measurement Results
|  | DC Voltage Test | [22 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  | MINIMUM CMC | FINAL UNCERTAINTY DECLARATION
|  | Range |  |  |  | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [6 empty] | Average |  | Error |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[13 empty] | 1 |  | 2 |  | 3 | [31 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  | 600 mV | [5 empty] | mV |  | - | [18 empty] | CoC | 0.00077 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.0016
[8 empty] | mV |  | - | [18 empty] | CoC | 0.0021 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.0018
[8 empty] | mV |  | - | [18 empty] | CoC | 0.0021 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.0018
|  | 6 V | [5 empty] | V |  | - | [18 empty] | CoC | 2.1e-06 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 1.8e-06
[8 empty] | V |  | - | [18 empty] | CoC | 1.8e-05 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 1e-05
[8 empty] | V |  | - | [18 empty] | CoC | 1.8e-05 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 1e-05
|  | 60 V | [5 empty] | V |  | - | [18 empty] | CoC | 1.8e-05 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 1e-05
[8 empty] | V |  | - | [18 empty] | CoC | 0.0001 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 1e-05
[8 empty] | V |  | - | [18 empty] | CoC | 0.00031 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.00018
[8 empty] | V |  | - | [18 empty] | CoC | 1.8e-05 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 1e-05
[8 empty] | V |  | - | [18 empty] | CoC | 0.00031 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.00018
|  | 600 V |  |  |  | 60 |  | V |  | - | [18 empty] | CoC | 0.00031 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.00018
[6 empty] | 540 |  | V |  | - | [18 empty] | CoC | 0.0033 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.0014
[6 empty] | -540 |  | V |  | - | [18 empty] | CoC | 0.0033 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.0014
|  | 1 000 V | [5 empty] | V |  | - | [18 empty] | CoC | 0.00031 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 1 |  |  |  | Repeatability | [8 empty] | 0.00018
[8 empty] | V |  | - | [18 empty] | CoC | 0.0033 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 1 |  |  |  | Repeatability | [8 empty] | 0.0014
[8 empty] | V |  | - | [18 empty] | CoC | 0.0033 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 1 |  |  |  | Repeatability | [8 empty] | 0.0014
|  | AC Voltage Test | [22 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  | MINIMUM CMC | FINAL UNCERTAINTY DECLARATION
|  | Range |  |  |  | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [6 empty] | Average |  | Error |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[13 empty] | 1 |  | 2 |  | 3 | [31 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  | 600 mV | [5 empty] | mV |  | 50 Hz | [18 empty] | CoC | 0.07 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.01
[8 empty] | mV |  | 1 kHz | [18 empty] | CoC | 0.07 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.01
[8 empty] | mV |  | 50 Hz | [18 empty] | CoC | 0.91 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.085
[8 empty] | mV |  | 1 kHz | [18 empty] | CoC | 0.91 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.085
|  | 6 V | [5 empty] | V |  | 50 Hz | [18 empty] | CoC | 0.00091 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 8.5e-05
[8 empty] | V |  | 1 kHz | [18 empty] | CoC | 0.00091 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 8.5e-05
[8 empty] | V |  | 50 Hz | [18 empty] | CoC | 0.0073 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 0.00085
[8 empty] | V |  | 1 kHz | [18 empty] | CoC | 0.0073 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 0.00085
|  | 60 V | [5 empty] | V |  | 50 Hz | [18 empty] | CoC | 0.0073 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.00085
[8 empty] | V |  | 1 kHz | [18 empty] | CoC | 0.0073 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.00085
[8 empty] | V |  | 50 Hz | [18 empty] | CoC | 0.0073 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.00085
[8 empty] | V |  | 1 kHz | [18 empty] | CoC | 0.0073 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.00085
[8 empty] | V |  | 50 Hz | [18 empty] | CoC | 0.071 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.0064
[8 empty] | V |  | 1 kHz | [18 empty] | CoC | 0.071 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.0064
|  | 600 V |  |  |  | 60 |  | V |  | 50 Hz | [18 empty] | CoC | 0.071 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.0064
[6 empty] | 60 |  | V |  | 1 kHz | [18 empty] | CoC | 0.071 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.0064
[6 empty] | 540 |  | V |  | 50 Hz | [18 empty] | CoC | 0.071 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.084
[6 empty] | 540 |  | V |  | 1 kHz | [18 empty] | CoC | 0.071 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.084
|  | 1 000 V | [5 empty] | V |  | 50 Hz | [18 empty] | CoC | 0.071 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 1 |  |  |  | Repeatability | [8 empty] | 0.0064
[8 empty] | V |  | 1 kHz | [18 empty] | CoC | 0.071 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 1 |  |  |  | Repeatability | [8 empty] | 0.0064
[8 empty] | V |  | 50 Hz | [18 empty] | CoC | 0.071 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 1 |  |  |  | Repeatability | [8 empty] | 0.084
[8 empty] | V |  | 1 kHz | [18 empty] | CoC | 0.071 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 1 |  |  |  | Repeatability | [8 empty] | 0.084
|  | Resistance Test | [22 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  | MINIMUM CMC | FINAL UNCERTAINTY DECLARATION
|  | Range |  |  |  | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [6 empty] | Average |  | Error |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[13 empty] | 1 |  | 2 |  | 3 | [31 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  | 600 Ω |  |  |  | 0 |  | Ω |  | - | [18 empty] | CoC | 7.7e-05 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.0078
[8 empty] | Ω |  | - | [18 empty] | CoC | 0.003 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.023
|  | 6 kΩ | [5 empty] | kΩ |  | - | [18 empty] | CoC | 5.5e-05 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 0.00023
|  | 60 kΩ | [5 empty] | kΩ |  | - | [18 empty] | CoC | 0.00081 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.0015
|  | 600 kΩ | [5 empty] | kΩ |  | - | [18 empty] | CoC | 0.0066 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.016
|  | 6 MΩ | [5 empty] | MΩ |  | - | [18 empty] | CoC | 0.00048 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 0.00053
|  | 50 MΩ | [5 empty] | MΩ |  | - | [18 empty] | CoC | 0.0047 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.015
|  | DC Current Test | [22 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  | MINIMUM CMC | FINAL UNCERTAINTY DECLARATION
|  | Range |  |  |  | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [6 empty] | Average |  | Error |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[13 empty] | 1 |  | 2 |  | 3 | [31 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  | 600 µA | [5 empty] | µA |  | - | [18 empty] | CoC | 0.09 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.064
|  | 6 000 µA | [5 empty] | µA |  | - | [18 empty] | CoC | 0.38 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 1 |  |  |  | Repeatability | [8 empty] | 0.45
|  | 60 mA | [5 empty] | mA |  | - | [18 empty] | CoC | 0.0014 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.0043
[8 empty] | mA |  | - | [18 empty] | CoC | 0.0014 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.0043
|  | 400 mA | [5 empty] | mA |  | - | [18 empty] | CoC | 0.027 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.082
|  | 6 A | [5 empty] | A |  | - | [18 empty] | CoC | 0.0014 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 0.0016
[8 empty] | A |  | - | [18 empty] | CoC | 0.0014 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 0.0016
|  | 10 A | [5 empty] | A |  | - | [18 empty] | CoC | 0.0014 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.0016
[8 empty] | A |  | - | [18 empty] | CoC | 0.0014 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.0016
|  | AC Current Test | [22 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  | MINIMUM CMC | FINAL UNCERTAINTY DECLARATION
|  | Range |  |  |  | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [6 empty] | Average |  | Error |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[13 empty] | 1 |  | 2 |  | 3 | [31 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  | 600 µA | [5 empty] | µA |  | 50 Hz | [18 empty] | CoC | 1.7 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.00037
[8 empty] | µA |  | 1 kHz | [18 empty] | CoC | 1.7 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.00037
|  | 6 000 µA | [5 empty] | µA |  | 50 Hz | [18 empty] | CoC | 13 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 1 |  |  |  | Repeatability | [8 empty] | 2.6
[8 empty] | µA |  | 1 kHz | [18 empty] | CoC | 13 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 1 |  |  |  | Repeatability | [8 empty] | 2.6
|  | 60 mA | [5 empty] | mA |  | 50 Hz | [18 empty] | CoC | 0.22 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.026
[8 empty] | mA |  | 1 kHz | [18 empty] | CoC | 0.22 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.026
|  | 400 mA | [5 empty] | mA |  | 50 Hz | [18 empty] | CoC | 0.72 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.21
[8 empty] | mA |  | 1 kHz | [18 empty] | CoC | 0.72 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.1 |  |  |  | Repeatability | [8 empty] | 0.21
|  | 6 A | [5 empty] | A |  | 50 Hz | [18 empty] | CoC | 0.01 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 0.0029
[8 empty] | A |  | 1 kHz | [18 empty] | CoC | 0.011 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.001 |  |  |  | Repeatability | [8 empty] | 0.0039
|  | 10 A | [5 empty] | A |  | 50 Hz | [18 empty] | CoC | 0.01 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.0029
[8 empty] | A |  | 1 kHz | [18 empty] | CoC | 0.011 |  | 2 |  | Annual Drift | 0 |  |  |  | Readability | 0.01 |  |  |  | Repeatability | [8 empty] | 0.0039
Calibrated By
Darwin F. Aliado
Calibration Officer's Name | [10 empty] | Signature | [7 empty] | Date
Checked By
Jennefer S. Ilas
Interim Manager (Electrical) | [10 empty] | Signature | [7 empty] | Date
Encoded By
- | [10 empty] | - | [7 empty] | -
Calibration Officer's Name | [10 empty] | Signature | [7 empty] | Date

## PASSFAIL (hidden)
PASS/FAIL ANALYSIS
Measurement Results |  |  | Based on Tolerance |  |  | Based on +/- Uncertainty Applied |  |  |  |  | FINAL REMARKS
Measured Value | Tolerance | Uncertainty | Upper Limit | Lower Limit | Remarks | Measured Value + Uncertainty | Upper Limit | Remarks | Measured Value - Uncertainty | Remarks
FINAL REMARKS
Measured Value | Tolerance | Uncertainty | Upper Limit | Lower Limit | Remarks | Measured + Upper UNC | Upper Limit | Remarks | Measured + Lower UNC | Remarks
FINAL REMARKS
Measured Value | Tolerance | Uncertainty | Upper Limit | Lower Limit | Remarks | Measured + Upper UNC | Upper Limit | Remarks | Measured + Lower UNC | Remarks
FINAL REMARKS
Measured Value | Tolerance | Uncertainty | Upper Limit | Lower Limit | Remarks | Measured + Upper UNC | Upper Limit | Remarks | Measured + Lower UNC | Remarks
FINAL REMARKS
Measured Value | Tolerance | Uncertainty | Upper Limit | Lower Limit | Remarks | Measured + Upper UNC | Upper Limit | Remarks | Measured + Lower UNC | Remarks