## Cal Certificate
`
Calibration Certificate
[11 empty] | Certificate No.: |  |  | 25- Rev.00
Customer : | [7 empty] | Location of Calibration :
| 0 | [7 empty] | SONJU ENGINEERING SERVICES
| 0 | [7 empty] | Electrical Calibration Area
Manufacturer: |  |  | Fluke |  |  |  |  | Description: |  |  |  | True RMS Digital Multimeter
Model No: |  |  | 289 |  |  |  |  | Options Installed: |  |  |  | none
Serial No: |  |  | 0 |  |  |  |  | Customer Asset No: |  |  |  | 0
Calibration Date: |  |  | 0 |  |  |  |  | Work Order No.: |  |  |  | 0
Received Date: |  |  | 0 |  |  |  |  | Customer P.O.: |  |  |  | none
Temperature: |  |  | 23 °C ± 5 °C |  |  |  |  | Humidity: |  |  |  | 55 %rh ± 15 %rh
Cal Method / Procedure :
As per Manufacturer's recommended performance tests procedure for Fluke 287/289 True RMS Digital Multimeter Calibration Manual, February 2008, Rev. 1, 3/09.
This is to certify that the above Measuring and Test Equipment has been calibrated by SONJU Engineering Services in accordance with relevant calibration procedure defined for this instrument.
Uncertainty:
The expanded uncertainty of measurement for the results given in this certificate was estimated with a coverage factor of k = 2 at 95% level of confidence. The uncertainties of measurement were calculated using a Budget Worksheet based on JCGM 100:2008 Evaluation of measurement data - Guide to the Expression of Uncertainty in Measurement, 1st edition- 2008.
Traceability:
The instrument was calibrated using reference standard/s traceable to SI Units as maintained by NML-ITDI (Philippines) through SONJU Engineering Services (Philippines).
Reference Standard Used:
| Description |  |  |  | Serial No. |  |  |  | Cal Report Ref. |  |  |  |  | Due Date
| Multiproduct Calibrator |  |  |  | 8100021 |  |  |  | 24-9973-01ALT Rev.00 |  |  |  |  | 2025-07-13
Remarks:
1. The results were the average of three (3) readings and were obtained during the time of calibration and refer only to the particular instrument submitted. Should the instrument be modified or damaged in any way, or develop inconsistent readings, the results may not be valid and the unit will require recalibration.
2. The user should determine the suitability of this instrument for its intended use.
3. Preliminary Inspection: No response at low current function.
4. Nominal values marked with " * " are outside scope of accreditation.
5. A positive error is to be subtracted to UUT indication and a negative error added to give true value.
6. As per customer request the next calibration due date is on: | [7 empty] | 1900-12-31
|  | 0 | [8 empty] | 0
|  | Calibration Officer | [8 empty] | Approved Signatory
Issue Date :
Measurement Results:
DC Voltage Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 50 mV |  | 0 mV |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 0.025 mV |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | -0.025 mV |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 50 mV |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 500 mV |  | 500 mV |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | -250 mV |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 5 V |  | 4 V |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 50 V |  | -40 V |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 500 V |  | 400 V |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 1 000 V |  | 600 V |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
DC/AC Voltage Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 500 mV |  | 50 mV |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 5 V |  | 0.2 V |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
AC/DC Voltage Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 500 mV |  | 250 mV / 35 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 5 V |  | 2 V / 5 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
AC Voltage Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 50 mV |  | 5 mV / 20 Hz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 50 mV / 65 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 500 mV |  | 50 mV / 100 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 250 mV / 65 Hz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 500 mV / 45 Hz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 5 V |  | 0.1 V / 60 Hz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 0.5 V / 10 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 3 V / 100 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 50 V |  | 15 V / 100 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 500 V |  | 500 V / 10 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 1 000 V |  | 1 000 V / 10 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
0
Calibration Officer
Measurement Results:
AC Voltage Frequency Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 500 mV |  | 45 Hz / 500 mV |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 950 kHz / 600 mV |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
AC Voltage Duty Cycle Test
|  | UUT Setting |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 5 V p-p Sq. Wave / 50 kHz |  | 15 % * |  | #DIV/0! |  |  |  | #DIV/0! |  |  | -
AC Voltage Peak Test
|  | UUT Setting |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 4 V p-p Sq. Wave, +1 V offset |  | 2 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
Temperature Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | Type K |  | 0 °C |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 100 °C |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 1 000.0 °C |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
Resistance Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 50 Ω |  | 0.2 Ω |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 50 Ω |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 500 Ω |  | 0 Ω |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 0.2 Ω |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 500 Ω |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 5 kΩ |  | 5 kΩ |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 50 kΩ |  | 50 kΩ |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 500 kΩ |  | 500 kΩ |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 5 MΩ |  | 5 MΩ |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 30 MΩ |  | 30 MΩ |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 500 MΩ |  | 300 MΩ |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
0
Calibration Officer
Measurement Results:
Capacitance Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 10 nF |  | 5 nF * |  | 5.10 nF |  |  |  | 0.10 nF |  |  | #DIV/0!
Diode Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 3.100 V |  | 3.5 kΩ |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
DC Curent Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 500 µA |  | 500 µA |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 5 000 µA |  | 5 000 µA |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 50 mA |  | 0.1 mA |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 50 mA |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 400 mA |  | 400 mA |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 5 A |  | 5 A |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 10 A |  | 10 A |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
AC Curent Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 5 00 µA |  | 500 µA / 60 Hz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 500 µA / 30 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 5 000 µA |  | 5 000 µA / 30 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 50 mA |  | 4 mA / 20 Hz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 30 mA / 30 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 400 mA |  | 300 mA / 30 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  |  |  | 400 mA / 60 Hz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 5 A |  | 5 A / 1 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
|  | 10 A |  | 5 A / 1 kHz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
LoZ Test
|  | Range |  | Nominal Value |  | Measured Value |  |  |  | Error |  |  | Uncertainty
|  | 1 000 V |  | 120 V / 60 Hz |  | #DIV/0! |  |  |  | #DIV/0! |  |  | #DIV/0!
0
Calibration Officer
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049)591-0807;
09171700815; 09171700231
Email: admin@sonju.ph
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049)591-0807;
09171700815; 09171700231
Email: admin@sonju.ph
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049)591-0807;
09171700815; 09171700231
Email: admin@sonju.ph
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049)591-0807;
09171700815; 09171700231
Email: admin@sonju.ph
Website: www.sonjuengineering.com

## DataSheet
CALIBRATION DATASHEET
|  |  | Work Order No.: | [18 empty] | Technical ID:
Unit Under Calibration
| Description | [5 empty] | : | True RMS Digital Multimeter | [8 empty] | Received Date | [5 empty] | :
| Manufacturer | [5 empty] | : | Fluke | [8 empty] | Calibration Date | [5 empty] | :
| Model | [5 empty] | : | 289 | [8 empty] | Options Installed | [5 empty] | : | none
| Serial Number | [5 empty] | : | [9 empty] | Customer P.O. | [5 empty] | : | none
| Range | [5 empty] | : | see specification sheet | [8 empty] | Asset Number | [5 empty] | :
| Res/Readability | [5 empty] | : | see specification sheet | [8 empty] | Accuracy | [5 empty] | : | see specification sheet
| Prev. SES Cal Cert: | [5 empty] | : | none | [8 empty] | Previous Technician | [5 empty] | : | none
| Customer Details | [15 empty] | Calibration Location
| Company |  |  | : | [14 empty] | x | In-House (SONJU Laboratory)
| Address |  |  | : | [15 empty] | On-Site Calibration (Please Specify)
Electrical Calibration Area
Reference Standard/s Used
| Description | [10 empty] | Serial No. |  |  |  |  | Cal Report Ref. | [8 empty] | Due Date
| Multiproduct Calibrator | [10 empty] | 8100021 |  |  |  |  | 24-9973-01ALT Rev.00 | [8 empty] | 2025-07-13
| - | [10 empty] | - |  |  |  |  | - | [8 empty] | -
Accessory Used
| Description | [10 empty] | Serial No. |  |  |  |  | Brand | [8 empty] | Model
| - | [10 empty] | - |  |  |  |  | - | [8 empty] | -
| - | [10 empty] | - |  |  |  |  | - | [8 empty] | -
| Environmental Condition | [15 empty] | Calibration Method
|  |  | Temperature | [6 empty] | Relative Humidity | [6 empty] | As per Manufacturer's recommended performance tests procedure for Fluke 287/289 True RMS Digital Multimeter Calibration Manual, February 2008, Rev. 1, 3/09.
[5 empty] | Start: |  |  | oC |  |  |  | Start: |  | %rh
[5 empty] | End: |  |  | oC |  |  |  | End: |  | %rh
Observations
Adjustment
Needed?
|  | - | [27 empty] | YES
|  | - | [26 empty] | x | NO
Measurement Results
|  |  | DC Voltage Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC (2024) |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 50 mV | [5 empty] | 0 |  | mV |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 0.02 | -0.02 | #DIV/0! | COC | 0.00031 | mV | 2 | 0.000155 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.001 | mV | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.00078 | mV | #DIV/0! | mV
[9 empty] | 0.025 |  | mV |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 0.045 | 0.005 | #DIV/0! | COC | 0.00077 | mV | 2 | 0.000385 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.001 | mV | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.00078 | mV | #DIV/0! | mV
[9 empty] | -0.025 |  | mV |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | -0.005 | -0.045 | #DIV/0! | COC | 0.00077 | mV | 2 | 0.000385 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.001 | mV | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.00078 | mV | #DIV/0! | mV
[9 empty] | 50 |  | mV |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 50.045 | 49.955 | #DIV/0! | COC | 0.00077 | mV | 2 | 0.000385 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.001 | mV | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.00078 | mV | #DIV/0! | mV
|  |  | 500 mV | [5 empty] | 500 |  | mV |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 500.15 | 499.85 | #DIV/0! | COC | 0.0021 | mV | 2 | 0.00105 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.01 | mV | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.0017 | mV | #DIV/0! | mV
[9 empty] | -250 |  | mV |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | -249.92 | -250.08 | #DIV/0! | COC | 0.00083 | mV | 2 | 0.000415 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.01 | mV | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.00078 | mV | #DIV/0! | mV
|  |  | 5 V | [5 empty] | 4 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 4.0012 | 3.9988 | #DIV/0! | COC | 1.8e-05 | V | 2 | 9e-06 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.0001 | V | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 1.7e-05 | V | #DIV/0! | V
|  |  | 50 V | [5 empty] | -40 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | -39.988 | -40.012 | #DIV/0! | COC | 0.00016 | V | 2 | 8e-05 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00013 | V | #DIV/0! | V
|  |  | 500 V | [5 empty] | 400 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 400.14 | 399.86 | #DIV/0! | COC | 0.0033 | V | 2 | 0.00165 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0013 | V | #DIV/0! | V
|  |  | 1 000 V | [5 empty] | 600 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 600.4 | 599.6 | #DIV/0! | COC | 0.0033 | V | 2 | 0.00165 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0013 | V | #DIV/0! | V
|  |  | DC/AC Voltage Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC (2024) |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
49.92 |  |  | 500 mV | [5 empty] | 50 |  | mV |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | COC | 0.00077 | mV | 2 | 0.000385 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.01 | mV | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.00078 | mV | #DIV/0! | mV
|  |  | 5 V | [5 empty] | 0.2 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | COC | 8.3e-07 | V | 2 | 4.15e-07 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.0001 | V | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 7.8e-07 | V | #DIV/0! | V
|  |  | AC/DC Voltage Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC (2024) |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 500 mV | [5 empty] | 250 |  | mV |  | 35 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 262.9 | 237.1 | #DIV/0! | COC | 0.07 | mV | 2 | 0.035 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.01 | mV | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.015 | mV | #DIV/0! | mV
|  |  | 5 V | [5 empty] | 2 |  | V |  | 5 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 2.036 | 1.964 | #DIV/0! | COC | 0.0015 | V | 2 | 0.00075 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.0001 | V | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 8.5e-05 | V | #DIV/0! | V
|  |  | AC Voltage Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC (2024) |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 50 mV | [5 empty] | 5 |  | mV |  | 20 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 5.135 | 4.865 | #DIV/0! | COC | 0.0088 | mV | 2 | 0.0044 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.001 | mV | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.0048 | mV | #DIV/0! | mV
[9 empty] | 50 |  | mV |  | 65 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 51.79 | 48.21 | #DIV/0! | COC | 0.59 | mV | 2 | 0.295 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.001 | mV | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.015 | mV | #DIV/0! | mV
|  |  | 500 mV | [5 empty] | 50 |  | mV |  | 100 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 52.15 | 47.85 | #DIV/0! | COC | 0.59 | mV | 2 | 0.295 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.01 | mV | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.045 | mV | #DIV/0! | mV
[9 empty] | 250 |  | mV |  | 65 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 259.15 | 240.85 | #DIV/0! | COC | 0.07 | mV | 2 | 0.035 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.01 | mV | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.01 | mV | #DIV/0! | mV
[9 empty] | 500 |  | mV |  | 45 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 501.75 | 498.25 | #DIV/0! | COC | 0.91 | mV | 2 | 0.455 | Std. Drift | 0 | mV | 1.7321 | 0 | Readability | 0.01 | mV | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.12 | mV | #DIV/0! | mV
|  |  | 5 V | [5 empty] | 0.1 |  | V |  | 60 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 0.1048 | 0.0952 | #DIV/0! | COC | 7e-05 | V | 2 | 3.5e-05 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.0001 | V | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 1e-05 | V | #DIV/0! | V
[9 empty] | 0.5 |  | V |  | 10 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 0.5055 | 0.4945 | #DIV/0! | COC | 0.0015 | V | 2 | 0.00075 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.0001 | V | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 8.5e-05 | V | #DIV/0! | V
[9 empty] | 3 |  | V |  | 100 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 3.184 | 2.816 | #DIV/0! | COC | 0.0059 | V | 2 | 0.00295 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.0001 | V | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00028 | V | #DIV/0! | V
|  |  | 50 V | [5 empty] | 15 |  | V |  | 100 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 15.565 | 14.435 | #DIV/0! | COC | 0.051 | V | 2 | 0.0255 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0035 | V | #DIV/0! | V
|  |  | 500 V | [5 empty] | 500 |  | V |  | 10 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 502.25 | 497.75 | #DIV/0! | COC | 0.1 | V | 2 | 0.05 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.084 | V | #DIV/0! | V
|  |  | 1 000 V | [5 empty] | 1000 |  | V |  | 10 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 1006.5 | 993.5 | #DIV/0! | COC | 0.1 | V | 2 | 0.05 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.084 | V | #DIV/0! | V
|  |  | AC Voltage Frequency Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Voltage |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 500 mV | [5 empty] | 45 |  | Hz |  | 500 mV | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 45.014 | 44.986 | #DIV/0! | COC | 5.8e-05 | Hz | 2 | 2.9e-05 | Std. Drift | 0 | Hz | 1.7321 | 0 | Readability | 0.001 | Hz | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | Hz | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Hz |  | Hz | #DIV/0! | Hz
|  |  | 500 mV | [5 empty] | 950 |  | kHz |  | 600 mV | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 950.1 | 949.9 | #DIV/0! | COC | 5.8e-05 | kHz | 2 | 2.9e-05 | Std. Drift | 0 | kHz | 1.7321 | 0 | Readability | 0.01 | kHz | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | kHz | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | kHz |  | kHz | #DIV/0! | kHz
AC Voltage Duty Cycle Test
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Standard Setting |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Result | Resolution
[16 empty] | 1 |  | 2 |  |  | 3
|  |  | 5 V | [5 empty] | 15 |  | % |  | 5 V p-p Sq. Wave / 50 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | 0 | 25.1 | 4.9 | #DIV/0! | 0.01
|  |  | AC Voltage Lo Pass Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 500 V | [5 empty] | 50 |  | V |  | 60 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 51.4 | 48.6 | #DIV/0! | COC | 0.071 | V | 2 | 0.0355 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V |  | V | #DIV/0! | V
|  |  | 500 V | [5 empty] | 50 |  | V |  | 1600 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 8 | 0 | #DIV/0! | COC | 0.1 | V | 2 | 0.05 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V |  | V | #DIV/0! | V
4
|  |  | AC Voltage Peak Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | UUT Setting |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Result | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 5 V | [5 empty] | 2 |  | kHz |  | 4 V p-p Sq. Wave, +1 V offset | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 2.137 | 1.863 | #DIV/0! | COC | 5.8e-05 | kHz | 2 | 2.9e-05 | Std. Drift | 0 | kHz | 1.7321 | 0 | Readability | 0.001 | kHz | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | kHz | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | kHz |  | kHz | #DIV/0! | kHz
Note: uV default unit symbol
|  |  | Temperature Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC (2024) |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | - | [5 empty] | 0 |  | °C |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 1 | -1 | #DIV/0! | COC | 0.00031 | mV | 2 | 1.55e-06 | Std. Drift | 0 | °C | 1.7321 | 0 | Readability | 0.1 | °C | 3.4641 | 0.028868 | Std. Dev of Reading | #DIV/0! | °C | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | °C | 0.4 | °C | #DIV/0! | °C
|  |  | - | [5 empty] | 100 |  | °C |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 102 | 98 | #DIV/0! | COC | 0.00077 | mV | 2 | 3.85e-06 | Std. Drift | 0 | °C | 1.7321 | 0 | Readability | 0.1 | °C | 3.4641 | 0.028868 | Std. Dev of Reading | #DIV/0! | °C | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | °C | 0.4 | °C | #DIV/0! | °C
|  |  | - | [5 empty] | 1000 |  | °C |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 1011 | 989 | #DIV/0! | COC | 0.00077 | mV | 2 | 3.85e-06 | Std. Drift | 0 | °C | 1.7321 | 0 | Readability | 0.1 | °C | 3.4641 | 0.028868 | Std. Dev of Reading | #DIV/0! | °C | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | °C | 0.4 | °C | #DIV/0! | °C
|  |  | Resistance Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC (2024) |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 50 Ω | [5 empty] | 0.2 |  | Ω |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 0.22 | 0.18 | #DIV/0! | COC | 7.1e-05 | Ω | 2 | 3.55e-05 | Std. Drift | 0 | Ω | 1.7321 | 0 | Readability | 0.001 | Ω | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | Ω | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Ω | 1.7e-06 | Ω | #DIV/0! | Ω
[9 empty] | 50 |  | Ω |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 50.095 | 49.905 | #DIV/0! | COC | 0.00052 | Ω | 2 | 0.00026 | Std. Drift | 0 | Ω | 1.7321 | 0 | Readability | 0.001 | Ω | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | Ω | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Ω | 0.012 | Ω | #DIV/0! | Ω
|  |  | 500 Ω | [5 empty] | 0 |  | Ω |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 0.1 | -0.1 | #DIV/0! | COC | 7.7e-05 | Ω | 2 | 3.85e-05 | Std. Drift | 0 | Ω | 1.7321 | 0 | Readability | 0.01 | Ω | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | Ω | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Ω | 1.7e-06 | Ω | #DIV/0! | Ω
[9 empty] | 0.2 |  | Ω |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 0.3 | 0.1 | #DIV/0! | COC | 7.1e-05 | Ω | 2 | 3.55e-05 | Std. Drift | 0 | Ω | 1.7321 | 0 | Readability | 0.01 | Ω | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | Ω | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Ω | 1.7e-06 | Ω | #DIV/0! | Ω
[9 empty] | 500 |  | Ω |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 500.35 | 499.65 | #DIV/0! | COC | 0.003 | Ω | 2 | 0.0015 | Std. Drift | 0 | Ω | 1.7321 | 0 | Readability | 0.01 | Ω | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | Ω | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Ω | 0.023 | Ω | #DIV/0! | Ω
|  |  | 5 kΩ | [5 empty] | 5 |  | kΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 5.0027 | 4.9973 | #DIV/0! | COC | 5.5e-05 | kΩ | 2 | 2.75e-05 | Std. Drift | 0 | kΩ | 1.7321 | 0 | Readability | 0.0001 | kΩ | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | kΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | kΩ | 0.00023 | kΩ | #DIV/0! | kΩ
|  |  | 50 kΩ | [5 empty] | 50 |  | kΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 50.027 | 49.973 | #DIV/0! | COC | 8.1e-05 | kΩ | 2 | 4.05e-05 | Std. Drift | 0 | kΩ | 1.7321 | 0 | Readability | 0.001 | kΩ | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | kΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | kΩ | 0.0015 | kΩ | #DIV/0! | kΩ
|  |  | 500 kΩ | [5 empty] | 500 |  | kΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 500.27 | 499.73 | #DIV/0! | COC | 0.0066 | kΩ | 2 | 0.0033 | Std. Drift | 0 | kΩ | 1.7321 | 0 | Readability | 0.01 | kΩ | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | kΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | kΩ | 0.016 | kΩ | #DIV/0! | kΩ
|  |  | 5 MΩ | [5 empty] | 5 |  | MΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 5.0079 | 4.9921 | #DIV/0! | COC | 0.00048 | MΩ | 2 | 0.00024 | Std. Drift | 0 | MΩ | 1.7321 | 0 | Readability | 0.0001 | MΩ | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | MΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | MΩ | 0.00053 | MΩ | #DIV/0! | MΩ
|  |  | 30 MΩ | [5 empty] | 30 |  | MΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 30.454 | 29.546 | #DIV/0! | COC | 0.0014 | MΩ | 2 | 0.0007 | Std. Drift | 0 | MΩ | 1.7321 | 0 | Readability | 0.001 | MΩ | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | MΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | MΩ | 0.0041 | MΩ | #DIV/0! | MΩ
|  |  | 500 MΩ | [5 empty] | 300 |  | MΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 324.2 | 275.8 | #DIV/0! | COC | 0.098 | MΩ | 2 | 0.049 | Std. Drift | 0 | MΩ | 1.7321 | 0 | Readability | 0.1 | MΩ | 3.4641 | 0.028868 | Std. Dev of Reading | #DIV/0! | MΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | MΩ | 0.33 | MΩ | #DIV/0! | MΩ
|  |  | Capacitance Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 10 |  |  | nF |  |  | 5 |  | nF |  | - | [11 empty] | 5.1 |  | 0.1 |  |  |  | 0.101 | 5.1 | 4.9 | #DIV/0! | SPECS | 0.0225 | nF | 2.58 | 0.0087209 | Std. Drift | 0 | nF | 1.7321 | 0 | Readability | 0.01 | nF | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | nF | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | nF | - | nF | #DIV/0! | nF
|  |  | Diode Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC (2024) |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
Range 5V |  |  | 3.100 V | [5 empty] | 3.5 |  | kΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | 0 | 3.1 | 2 | #DIV/0! | COC | 5.5e-05 | kΩ | 2 | 2.75e-05 | Std. Drift | 0 | kΩ | 1.7321 | 0 | Readability | 0.0001 | kΩ | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | kΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | kΩ | 0.00023 | kΩ | #DIV/0! | kΩ
|  |  | DC Curent Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC (2024) |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 500 µA | [5 empty] | 500 |  | µA |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 500.58 | 499.42 | #DIV/0! | COC | 0.09 | µA | 2 | 0.045 | Std. Drift | 0 | µA | 1.7321 | 0 | Readability | 0.01 | µA | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | µA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | µA | 0.064 | µA | #DIV/0! | µA
|  |  | 5 000 µA | [5 empty] | 5000 |  | µA |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 5004 | 4996 | #DIV/0! | COC | 0.38 | µA | 2 | 0.19 | Std. Drift | 0 | µA | 1.7321 | 0 | Readability | 0.1 | µA | 3.4641 | 0.028868 | Std. Dev of Reading | #DIV/0! | µA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | µA | 0.45 | µA | #DIV/0! | µA
|  |  | 50 mA | [5 empty] | 0.1 |  | mA |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 0.11 | 0.09 | #DIV/0! | COC | 8.2e-06 | mA | 2 | 4.1e-06 | Std. Drift | 0 | mA | 1.7321 | 0 | Readability | 0.001 | mA | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 1.6e-05 | mA | #DIV/0! | mA
[9 empty] | 50 |  | mA |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 50.035 | 49.965 | #DIV/0! | COC | 0.0014 | mA | 2 | 0.0007 | Std. Drift | 0 | mA | 1.7321 | 0 | Readability | 0.001 | mA | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.0043 | mA | #DIV/0! | mA
|  |  | 400 mA | [5 empty] | 400 |  | mA |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 400.62 | 399.38 | #DIV/0! | COC | 0.027 | mA | 2 | 0.0135 | Std. Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.082 | mA | #DIV/0! | mA
|  |  | 5 A | [5 empty] | 5 |  | A |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 5.016 | 4.984 | #DIV/0! | COC | 0.0014 | A | 2 | 0.0007 | Std. Drift | 0 | A | 1.7321 | 0 | Readability | 0.0001 | A | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | A | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | A | 0.0016 | A | #DIV/0! | A
|  |  | 10 A | [5 empty] | 10 |  | A |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 10.032 | 9.968 | #DIV/0! | COC | 0.0014 | A | 2 | 0.0007 | Std. Drift | 0 | A | 1.7321 | 0 | Readability | 0.001 | A | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | A | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | A | 0.0016 | A | #DIV/0! | A
|  |  | AC Curent Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC (2024) |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 5 00 µA | [5 empty] | 500 |  | µA |  | 60 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 503.2 | 496.8 | #DIV/0! | COC | 1.7 | µA | 2 | 0.85 | Std. Drift | 0 | µA | 1.7321 | 0 | Readability | 0.01 | µA | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | µA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | µA | 0.37 | µA | #DIV/0! | µA
[9 empty] | 500 |  | µA |  | 30 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 507.15 | 492.85 | #DIV/0! | COC | 1.7 | µA | 2 | 0.85 | Std. Drift | 0 | µA | 1.7321 | 0 | Readability | 0.01 | µA | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | µA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | µA | 3 | µA | #DIV/0! | µA
|  |  | 5 000 µA | [5 empty] | 5000 |  | µA |  | 30 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 5071.5 | 4928.5 | #DIV/0! | COC | 13 | µA | 2 | 6.5 | Std. Drift | 0 | µA | 1.7321 | 0 | Readability | 0.1 | µA | 3.4641 | 0.028868 | Std. Dev of Reading | #DIV/0! | µA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | µA | 13 | µA | #DIV/0! | µA
|  |  | 50 mA | [5 empty] | 4 |  | mA |  | 20 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 4.06 | 3.94 | #DIV/0! | COC | 0.013 | mA | 2 | 0.0065 | Std. Drift | 0 | mA | 1.7321 | 0 | Readability | 0.001 | mA | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.0062 | mA | #DIV/0! | mA
[9 empty] | 30 |  | mA |  | 30 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 30.625 | 29.375 | #DIV/0! | COC | 0.031 | mA | 2 | 0.0155 | Std. Drift | 0 | mA | 1.7321 | 0 | Readability | 0.001 | mA | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.013 | mA | #DIV/0! | mA
|  |  | 400 mA | [5 empty] | 300 |  | mA |  | 30 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 315.4 | 284.6 | #DIV/0! | COC | 0.0013 | mA | 2 | 0.00065 | Std. Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.26 | mA | #DIV/0! | mA
[9 empty] | 400 |  | mA |  | 60 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 402.45 | 397.55 | #DIV/0! | COC | 0.72 | mA | 2 | 0.36 | Std. Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Std. Dev of Reading | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.21 | mA | #DIV/0! | mA
|  |  | 5 A | [5 empty] | 5 |  | A |  | 1 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 5.042 | 4.958 | #DIV/0! | COC | 0.011 | A | 2 | 0.0055 | Std. Drift | 0 | A | 1.7321 | 0 | Readability | 0.0001 | A | 3.4641 | 2.8868e-05 | Std. Dev of Reading | #DIV/0! | A | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | A | 0.0039 | A | #DIV/0! | A
|  |  | 10 A | [5 empty] | 5 |  | A |  | 1 kHz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 5.045 | 4.955 | #DIV/0! | COC | 0.011 | A | 2 | 0.0055 | Std. Drift | 0 | A | 1.7321 | 0 | Readability | 0.001 | A | 3.4641 | 0.00028868 | Std. Dev of Reading | #DIV/0! | A | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | A | 0.0039 | A | #DIV/0! | A
|  |  | LoZ Test | [31 empty] | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC (2024) |  | FINAL UNCERTAINTY DECLARATION
|  |  | Range | [5 empty] | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Specs Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | Source | Uncertainty |  | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[16 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
|  |  | 1 000 V | [5 empty] | 120 |  | V |  | 60 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | 126.4 | 113.6 | #DIV/0! | COC | 0.071 | V | 2 | 0.0355 | Std. Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Std. Dev of Reading | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0064 | V | #DIV/0! | V
Calibrated By
0
| Calibration Officer's Name | [13 empty] | Signature | [7 empty] | Date
Checked By
2025-06-26
| Interim Manager (Electrical) | [13 empty] | Signature | [7 empty] | Date
Encoded By
| - | [21 empty] | -
| Calibration Officer's Name | [13 empty] | Signature | [7 empty] | Date