## Cal Certificate
Calibration Certificate
[10 empty] | Certificate No.: |  |  |  | 25- Rev.00
Customer : | [6 empty] | Location of Calibration :
| 0 | [6 empty] | SONJU ENGINEERING SERVICES
| 0 | [6 empty] | Electrical Calibration Area
Manufacturer: |  |  | Sanwa |  |  |  | Description: |  |  |  |  | Digital Multimeter
Model No: |  |  | CD800a |  |  |  | Options Installed: |  |  |  |  | none
Serial No: |  |  | 0 |  |  |  | Customer Asset No: |  |  |  |  | 0
Calibration Date: |  |  | 0 |  |  |  | Work Order No.: |  |  |  |  | 0
Received Date: |  |  | 0 |  |  |  | Customer P.O.: |  |  |  |  | none
Temperature: |  |  | 23 °C ± 5 °C |  |  |  | Humidity: |  |  |  |  | 55 %rh ± 15 %rh
Cal Method / Procedure :
As per SONJU Engineering Services Work Instruction 7W-02-54 Digital Multimeter Calibration Procedure.
This is to certify that the above Measuring and Test Equipment has been calibrated by SONJU Engineering Services in accordance with relevant calibration procedure defined for this instrument.
Uncertainty:
The expanded uncertainty of measurement for the results given in this certificate was estimated with a coverage factor of k = 2 at 95% level of confidence. The uncertainties of measurement were calculated using a budget worksheet based on JCGM 100:2008 Evaluation of measurement data - Guide to the Expression of Uncertainty in Measurement, 1st edition- 2008.
Traceability:
The instrument was calibrated using reference standard/s traceable to SI Units as maintained by NPL (United Kingdom), ITC-SCL (Hong Kong) and NML-ITDI (Philippines) through SONJU Engineering Services (Philippines).
Reference Standard Used:
| Description |  |  |  | Serial No. |  |  | Cal Report Ref. | [5 empty] | Due Date
| Multiproduct Calibrator |  |  |  | 2181901 |  |  | 25-1604-01JSC Rev.00 | [5 empty] | 2026-02-07
Remarks:
1. The results were the average of the three (3) readings and were obtained during the time of calibration and refer only to the particular instrument submitted. Should the instrument be modified or damaged in any way, or develop inconsistent readings, the results may not be valid and the unit will require recalibration.
2. The user should determine the suitability of this instrument for its intended use.
3. Preliminary Inspection: Satisfactory.
4. A positve error is to subtracted to the UUT indication and negative error added to give true value.
5. As per customer request the next calibration due date is on: | [6 empty] | 1900-12-31
|  | 0 | [8 empty] | Jennefer S. Ilas
|  | Calibration Officer | [8 empty] | Approved Signatory
Issue Date :
Measurement Results:
DC Voltage Test
|  | Range |  | Nominal Value |  | Measured Value |  |  | Error |  |  |  | Uncertainty
|  | 400 mV |  | 40 mV |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 360 mV |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | -360 mV |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 4 V |  | 0.4 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 3.6 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | -3.6 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 40 V |  | 4 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 20 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 36 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | -4 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | -36 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 400 V |  | 40 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 360 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | -360 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 600 V |  | 60 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 540 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | -540 V |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
Resistance Test
|  | Range |  | Nominal Value |  | Measured Value |  |  | Error |  |  |  | Uncertainty
|  | 400 Ω |  | 0 Ω |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 360 Ω |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 4 kΩ |  | 3.6 kΩ |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 40 kΩ |  | 36 kΩ |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 400 kΩ |  | 360 kΩ |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 4 MΩ |  | 3.6 MΩ |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 40 MΩ |  | 36 MΩ |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
0
Calibration Officer
Measurement Results:
AC Voltage Test
|  | Range |  | Nominal Value |  | Measured Value |  |  | Error |  |  |  | Uncertainty
|  | 4 V |  | 0.4 V / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 0.4 V / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 3.6 V / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 3.6 V / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 40 V |  | 4 V / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 4 V / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 20 V / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 20 V / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 36 V / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 36 V / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 400 V |  | 40 V / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 40 V / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 360 V / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 360 V / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 600 V |  | 60 V / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 60 V / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 540 V / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 540 V / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
DC Current Test
|  | Range |  | Nominal Value |  | Measured Value |  |  | Error |  |  |  | Uncertainty
|  | 40 mA |  | 36 mA |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | -36 mA |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 400 mA |  | 360 mA |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
AC Current Test
|  | Range |  | Nominal Value |  | Measured Value |  |  | Error |  |  |  | Uncertainty
|  | 40 mA |  | 36 mA / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 36 mA / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  | 400 mA |  | 360 mA / 50 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
|  |  |  | 360 mA / 400 Hz |  | #DIV/0! |  |  | #DIV/0! |  |  |  | #DIV/0!
0
Calibration Officer
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049)591-0807;
09171700815; 09171700231
Email: admin@sonju.ph
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049)591-0807;
09171700815; 09171700231
Email: admin@sonju.ph
Website: www.sonjuengineering.com
This report shall not be reproduced except in full,
without the written approval from the laboratory
Block 1 Lot 49, Jubilation Commercial Strip,
Brgy. Zapote, Biñan City, Laguna 4024 Phils.
Contact Nos.: (049)591-0807;
09171700815; 09171700231
Email: admin@sonju.ph
Website: www.sonjuengineering.com

## DataSheet
CALIBRATION DATASHEET
| Work Order No.: | [17 empty] | Technical ID:
Unit Under Calibration
Description |  |  |  | : | Digital Multimeter | [8 empty] | Received Date | [5 empty] | :
Manufacturer |  |  |  | : | Sanwa | [8 empty] | Calibration Date | [5 empty] | :
Model |  |  |  | : | CD800a | [8 empty] | Options Installed | [5 empty] | : | none
Serial Number |  |  |  | : | [9 empty] | Customer P.O. | [5 empty] | : | none
Range |  |  |  | : | see specification sheet | [8 empty] | Asset Number | [5 empty] | :
Res/Readability |  |  |  | : | 4000 counts | [8 empty] | Accuracy | [5 empty] | : | see specification sheet
Prev. SES Cal Cert: |  |  |  | : | [9 empty] | Previous Technician | [5 empty] | :
Customer Details | [13 empty] | Calibration Location
Company |  | : | [13 empty] | x | In-House (SONJU Laboratory)
Address |  | : | [14 empty] | On-Site Calibration (Please Specify)
Electrical Calibration Area
Reference Standard/s Used
Description | [8 empty] | Serial No. |  |  |  |  | Cal Report Ref. | [8 empty] | Due Date
Multiproduct Calibrator | [8 empty] | 2181901 |  |  |  |  | 25-1604-01JSC Rev.00 | [8 empty] | 2026-02-07
Accessory Used
Description | [8 empty] | Serial No. |  |  |  |  | Brand | [8 empty] | Model
- | [8 empty] | - |  |  |  |  | - | [8 empty] | -
Environmental Condition | [13 empty] | Calibration Method
|  |  | Temperature | [5 empty] | Relative Humidity |  |  |  |  | As per SONJU Engineering Services Work Instruction 7W-02-54 Digital Multimeter Calibration Procedure.
|  |  |  | Start: |  |  | oC |  |  | Start: |  | %rh
|  |  |  | End: |  |  | oC |  |  | End: |  | %rh
Observations
Adjustment
Needed?
YES
[26 empty] | x | NO
Measurement Results
| DC Voltage Test | [26 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
| Range |  |  |  |  | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[13 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
| 400 mV |  |  |  |  | 40 |  | mV |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.00082 | mV | 2 | 0.00041 | Annual Drift | 0 | mV | 1.7321 | 0 | Readability | 0.1 | mV | 3.4641 | 0.028868 | Repeatability | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.0016 | mV | #DIV/0! | mV
[6 empty] | 360 |  | mV |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0024 | mV | 2 | 0.0012 | Annual Drift | 0 | mV | 1.7321 | 0 | Readability | 0.1 | mV | 3.4641 | 0.028868 | Repeatability | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.0018 | mV | #DIV/0! | mV
[6 empty] | -360 |  | mV |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.013 | mV | 2 | 0.0065 | Annual Drift | 0 | mV | 1.7321 | 0 | Readability | 0.1 | mV | 3.4641 | 0.028868 | Repeatability | #DIV/0! | mV | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mV | 0.0018 | mV | #DIV/0! | mV
| 4 V |  |  |  |  | 0.4 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 2.4e-06 | V | 2 | 1.2e-06 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 1.8e-06 | V | #DIV/0! | V
[6 empty] | 3.6 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 2.1e-05 | V | 2 | 1.05e-05 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 1e-05 | V | #DIV/0! | V
[6 empty] | -3.6 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 2.2e-05 | V | 2 | 1.1e-05 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 1e-05 | V | #DIV/0! | V
| 40 V |  |  |  |  | 4 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 2.1e-05 | V | 2 | 1.05e-05 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 1e-05 | V | #DIV/0! | V
[6 empty] | 20 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0004 | V | 2 | 0.0002 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 1e-05 | V | #DIV/0! | V
[6 empty] | 36 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0004 | V | 2 | 0.0002 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00018 | V | #DIV/0! | V
[6 empty] | -4 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 2.2e-05 | V | 2 | 1.1e-05 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 1e-05 | V | #DIV/0! | V
[6 empty] | -36 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0004 | V | 2 | 0.0002 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00018 | V | #DIV/0! | V
| 400 V |  |  |  |  | 40 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0004 | V | 2 | 0.0002 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00018 | V | #DIV/0! | V
[6 empty] | 360 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0047 | V | 2 | 0.00235 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0014 | V | #DIV/0! | V
[6 empty] | -360 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0047 | V | 2 | 0.00235 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0014 | V | #DIV/0! | V
| 600 V |  |  |  |  | 60 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.00041 | V | 2 | 0.000205 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00018 | V | #DIV/0! | V
[6 empty] | 540 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0047 | V | 2 | 0.00235 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0014 | V | #DIV/0! | V
[6 empty] | -540 |  | V |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0047 | V | 2 | 0.00235 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0014 | V | #DIV/0! | V
| AC Voltage Test | [26 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
| Range |  |  |  |  | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[13 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
| 4 V |  |  |  |  | 0.4 |  | V |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.00091 | V | 2 | 0.000455 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 8.5e-05 | V | #DIV/0! | V
[6 empty] | 0.4 |  | V |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.00091 | V | 2 | 0.000455 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 8.5e-05 | V | #DIV/0! | V
[6 empty] | 3.6 |  | V |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00085 | V | #DIV/0! | V
[6 empty] | 3.6 |  | V |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.001 | V | 3.4641 | 0.00028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00085 | V | #DIV/0! | V
| 40 V |  |  |  |  | 4 |  | V |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00085 | V | #DIV/0! | V
[6 empty] | 4 |  | V |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00085 | V | #DIV/0! | V
[6 empty] | 20 |  | V |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00085 | V | #DIV/0! | V
[6 empty] | 20 |  | V |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0073 | V | 2 | 0.00365 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.00085 | V | #DIV/0! | V
[6 empty] | 36 |  | V |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0064 | V | #DIV/0! | V
[6 empty] | 36 |  | V |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.01 | V | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0064 | V | #DIV/0! | V
| 400 V |  |  |  |  | 40 |  | V |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0064 | V | #DIV/0! | V
[6 empty] | 40 |  | V |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0064 | V | #DIV/0! | V
[6 empty] | 360 |  | V |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.084 | V | #DIV/0! | V
[6 empty] | 360 |  | V |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 0.1 | V | 3.4641 | 0.028868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.084 | V | #DIV/0! | V
| 600 V |  |  |  |  | 60 |  | V |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0064 | V | #DIV/0! | V
[6 empty] | 60 |  | V |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.0064 | V | #DIV/0! | V
[6 empty] | 540 |  | V |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.084 | V | #DIV/0! | V
[6 empty] | 540 |  | V |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.071 | V | 2 | 0.0355 | Annual Drift | 0 | V | 1.7321 | 0 | Readability | 1 | V | 3.4641 | 0.28868 | Repeatability | #DIV/0! | V | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | V | 0.084 | V | #DIV/0! | V
| Resistance Test | [26 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
| Range |  |  |  |  | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[13 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
| 400 Ω |  |  |  |  | 0 |  | Ω |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 5.8e-05 | Ω | 2 | 2.9e-05 | Annual Drift | 0 | Ω | 1.7321 | 0 | Readability | 0.1 | Ω | 3.4641 | 0.028868 | Repeatability | #DIV/0! | Ω | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Ω | 0.0078 | Ω | #DIV/0! | Ω
[6 empty] | 360 |  | Ω |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0013 | Ω | 2 | 0.00065 | Annual Drift | 0 | Ω | 1.7321 | 0 | Readability | 0.1 | Ω | 3.4641 | 0.028868 | Repeatability | #DIV/0! | Ω | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Ω | 0.023 | Ω | #DIV/0! | Ω
| 4 kΩ |  |  |  |  | 3.6 |  | kΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 9.4e-06 | kΩ | 2 | 4.7e-06 | Annual Drift | 0 | kΩ | 1.7321 | 0 | Readability | 0.001 | kΩ | 3.4641 | 0.00028868 | Repeatability | #DIV/0! | kΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | kΩ | 0.00023 | kΩ | #DIV/0! | kΩ
| 40 kΩ |  |  |  |  | 36 |  | kΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.00044 | kΩ | 2 | 0.00022 | Annual Drift | 0 | kΩ | 1.7321 | 0 | Readability | 0.01 | kΩ | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | kΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | kΩ | 0.0015 | kΩ | #DIV/0! | kΩ
| 400 kΩ |  |  |  |  | 360 |  | kΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.11 | kΩ | 2 | 0.055 | Annual Drift | 0 | kΩ | 1.7321 | 0 | Readability | 0.1 | kΩ | 3.4641 | 0.028868 | Repeatability | #DIV/0! | kΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | kΩ | 0.016 | kΩ | #DIV/0! | kΩ
| 4 MΩ |  |  |  |  | 3.6 |  | MΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.00011 | MΩ | 2 | 5.5e-05 | Annual Drift | 0 | MΩ | 1.7321 | 0 | Readability | 0.001 | MΩ | 3.4641 | 0.00028868 | Repeatability | #DIV/0! | MΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | MΩ | 0.00053 | MΩ | #DIV/0! | MΩ
| 40 MΩ |  |  |  |  | 36 |  | MΩ |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.0034 | MΩ | 2 | 0.0017 | Annual Drift | 0 | MΩ | 1.7321 | 0 | Readability | 0.01 | MΩ | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | MΩ | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | MΩ | 0.015 | MΩ | #DIV/0! | MΩ
| DC Current Test | [26 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
| Range |  |  |  |  | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[13 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
| 40 mA |  |  |  |  | 36 |  | mA |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.00083 | mA | 2 | 0.000415 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.0043 | mA | #DIV/0! | mA
[6 empty] | -36 |  | mA |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.00083 | mA | 2 | 0.000415 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.0043 | mA | #DIV/0! | mA
| 400 mA |  |  |  |  | 360 |  | mA |  | - | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.014 | mA | 2 | 0.007 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.1 | mA | 3.4641 | 0.028868 | Repeatability | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.082 | mA | #DIV/0! | mA
| AC Current Test | [26 empty] | Based on Tolerance |  |  |  | Standard |  |  |  |  | Standard |  |  |  |  | Unit Under Test |  |  |  |  | Unit Under Test |  |  |  |  | Expanded Uncertainty |  |  |  |  | MINIMUM CMC |  | FINAL UNCERTAINTY DECLARATION
| Range |  |  |  |  | Nominal Value |  |  |  | Frequency |  |  | Measured Value | [8 empty] | Average |  | Error |  |  |  | Tolerance | Upper Limit | Lower Limit | Remarks | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | Source | Uncertainty | Unit | Divisor | Ui | COMBINED STANDARD | EFFECTIVE DEGREES | C. FACTOR, k = student's | EXPANDED UNCERTAINTY,
[13 empty] | 1 |  | 2 |  |  | 3 | [33 empty] | UNCERTAINTY, Uc | OF FREEDOM, Veff | t for Veff and 95 % C.L. | U @ 95 % C.L. ±
| 40 mA |  |  |  |  | 36 |  | mA |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.22 | mA | 2 | 0.11 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.026 | mA | #DIV/0! | mA
[6 empty] | 36 |  | mA |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 0.22 | mA | 2 | 0.11 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.01 | mA | 3.4641 | 0.0028868 | Repeatability | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.026 | mA | #DIV/0! | mA
| 400 mA |  |  |  |  | 360 |  | mA |  | 50 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 2 | mA | 2 | 1 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.1 | mA | 3.4641 | 0.028868 | Repeatability | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.21 | mA | #DIV/0! | mA
[6 empty] | 360 |  | mA |  | 400 Hz | [11 empty] | #DIV/0! |  | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | CoC | 2 | mA | 2 | 1 | Annual Drift | 0 | mA | 1.7321 | 0 | Readability | 0.1 | mA | 3.4641 | 0.028868 | Repeatability | #DIV/0! | mA | 1.7321 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | mA | 0.21 | mA | #DIV/0! | mA
Calibrated By
0
Calibration Officer's Name | [10 empty] | Signature | [8 empty] | Date
Checked By
Interim Manager (Electrical) | [10 empty] | Signature | [8 empty] | Date
Encoded By
- | [10 empty] | - | [8 empty] | -
Calibration Officer's Name | [10 empty] | Signature | [8 empty] | Date

## PASSFAIL (hidden)
PASS/FAIL ANALYSIS
Measurement Results | [5 empty] | Based on Tolerance |  |  |  |  | Based on +/- Uncertainty Applied | [7 empty] | FINAL REMARKS
DC Voltage Test
Measured Value |  | Tolerance |  | Uncertainty |  | Upper Limit |  | Lower Limit |  | Remarks | Measured Value + Uncertainty |  | Upper Limit |  | Remarks | Measured Value - Uncertainty |  | Remarks
#DIV/0! | mV | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | #DIV/0! | mV | #DIV/0! | #DIV/0!
#DIV/0! | mV | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | #DIV/0! | mV | #DIV/0! | #DIV/0!
#DIV/0! | mV | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | #DIV/0! | mV | #DIV/0! | mV | #DIV/0! | #DIV/0! | mV | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
AC Voltage Test | [18 empty] | FINAL REMARKS
Measured Value |  | Tolerance |  | Uncertainty |  | Upper Limit |  | Lower Limit |  | Remarks | Measured + Upper UNC |  | Upper Limit |  | Remarks | Measured + Lower UNC |  | Remarks
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
#DIV/0! | V | #DIV/0! | V | #DIV/0! | V | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | V | #DIV/0! | V | #DIV/0! | #DIV/0! | V | #DIV/0! | #DIV/0!
Resistance Test | [18 empty] | FINAL REMARKS
Measured Value |  | Tolerance |  | Uncertainty |  | Upper Limit |  | Lower Limit |  | Remarks | Measured + Upper UNC |  | Upper Limit |  | Remarks | Measured + Lower UNC |  | Remarks
#DIV/0! | Ω | #DIV/0! | Ω | #DIV/0! | Ω | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | Ω | #DIV/0! | Ω | #DIV/0! | #DIV/0! | Ω | #DIV/0! | #DIV/0!
#DIV/0! | Ω | #DIV/0! | Ω | #DIV/0! | Ω | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | Ω | #DIV/0! | Ω | #DIV/0! | #DIV/0! | Ω | #DIV/0! | #DIV/0!
#DIV/0! | kΩ | #DIV/0! | kΩ | #DIV/0! | kΩ | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | kΩ | #DIV/0! | kΩ | #DIV/0! | #DIV/0! | kΩ | #DIV/0! | #DIV/0!
#DIV/0! | kΩ | #DIV/0! | kΩ | #DIV/0! | kΩ | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | kΩ | #DIV/0! | kΩ | #DIV/0! | #DIV/0! | kΩ | #DIV/0! | #DIV/0!
#DIV/0! | kΩ | #DIV/0! | kΩ | #DIV/0! | kΩ | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | kΩ | #DIV/0! | kΩ | #DIV/0! | #DIV/0! | kΩ | #DIV/0! | #DIV/0!
#DIV/0! | MΩ | #DIV/0! | MΩ | #DIV/0! | MΩ | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | MΩ | #DIV/0! | MΩ | #DIV/0! | #DIV/0! | MΩ | #DIV/0! | #DIV/0!
#DIV/0! | MΩ | #DIV/0! | MΩ | #DIV/0! | MΩ | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | MΩ | #DIV/0! | MΩ | #DIV/0! | #DIV/0! | MΩ | #DIV/0! | #DIV/0!
DC Current Test | [18 empty] | FINAL REMARKS
Measured Value |  | Tolerance |  | Uncertainty |  | Upper Limit |  | Lower Limit |  | Remarks | Measured + Upper UNC |  | Upper Limit |  | Remarks | Measured + Lower UNC |  | Remarks
#DIV/0! | mA | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! | #DIV/0! | mA | #DIV/0! | #DIV/0!
#DIV/0! | mA | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! | #DIV/0! | mA | #DIV/0! | #DIV/0!
#DIV/0! | mA | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! | #DIV/0! | mA | #DIV/0! | #DIV/0!
AC Current Test | [18 empty] | FINAL REMARKS
Measured Value |  | Tolerance |  | Uncertainty |  | Upper Limit |  | Lower Limit |  | Remarks | Measured + Upper UNC |  | Upper Limit |  | Remarks | Measured + Lower UNC |  | Remarks
#DIV/0! | mA | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! | #DIV/0! | mA | #DIV/0! | #DIV/0!
#DIV/0! | mA | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! | #DIV/0! | mA | #DIV/0! | #DIV/0!
#DIV/0! | mA | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! | #DIV/0! | mA | #DIV/0! | #DIV/0!
#DIV/0! | mA | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! |  | #DIV/0! |  | #DIV/0! | #DIV/0! | mA | #DIV/0! | mA | #DIV/0! | #DIV/0! | mA | #DIV/0! | #DIV/0!